## SLAC Traveler for HE-SXR Undulator #29A804-000 in MMF

Dataset 003

This traveler is intended to cover preparation, magnetic measurements, and fiducialization of the SXR undulator segments at MMF.

Preparation:

- Remove all protective covers. Install wooden platform between the undulator and granite.

- Remove plastic cover. Clean up the granite.

- Attach temperature sensors.

- Attach ground wire for capacitive sensors.

- Check linear potentiometers. Install brackets if necessary.

- Label modules from 1 to 3.

- Install coil supports.

| Vendor serial number from undulator label: |  |
| --- | --- |
| Technician (initials): |  |
| Date/Time: |  |

- Call controls group to hook-up and check the control system. Check completed.

- Call alignment crew to align the undulator to the bench. Alignment done.

Mechanical measurements:

- Open gap to 10mm (tuning gap). Check gap straightness with the touch probe.

- Correct encoders’ offsets to make jaws parallel and level.

- Set half-gap encoders’ offsets at 10mm gap to read 5.0000mm.

- Measure every pole at 10mm gap, correct pole positions to ±20µm, if necessary.

- At 10mm gap set the midplane encoders’ offsets, so the encoders read 5.0000mm.

- Measure mechanical axis shift vs. Gap; gaps = 7.5,8.5,10,13,17,25, and 33mm. Make sure it stays close to zero when corrections applied, otherwise call controls group to fix the issue.

- Check undulator pitch vs Gap with the touch probe, gaps = 7.5,8.5,10,13,17,25 and 33mm. Make sure it is close to zero, otherwise call controls group to fix the issue.

- Remove touch probe. Install the Hall probe, align.

- Find the third pole z-position.                                  Z3 =

| Engineer(initials)/date and time: |  |
| --- | --- |

Tuning.

Follow the check list below, consult with the tuning test plan (LCLS-TN-22-??).

Tuning process checklist:

| Engineer (initials): |  |
| --- | --- |
| Tuning started, Date (mm-dd-yyyy): |  |
| Undulator magnetically aligned at tuning Gap = 10mm (adjust encoder offsets if necessary), results recorded into the table below. |
| MA shift vs Gap measured with Hall probe for gaps = 7.2,8.5,10,13,17,25, and 33 mm |
| Fields measured for central tuning x = ± 2mm at tuning gap=10mm |
| Upstream PM measured |
| Capacitive sensor measurements before tuning done. Gap = 7.4mm. |
| Kugler bench OFF, Long coil set, PDI and XPS ON |
| Field integrals measured for tuning trajectories |
| Coil and stages in park position, PDI and XPS OFF, the coil removed for shimming. |
| Shim table prepared. Gap opened to 180mm. |
| Undulator trajectory shimming completed |
| Kugler bench ON and re-initialized |
| Capacitive sensor measurements after tuning done |
| Trajectories and phase errors checked. (If additional shimming needed, repeat Hall probe measurements, coil measurements and tuning). |
| Trajectories and phase errors tuned to tolerances. |
| Fields measured for tuning ends. |
| Kugler bench OFF. Long coil set, PDI and XPS ON |
| Field integrals measured for tuning first and second field integrals. |
| Shim table prepared. Gap opened to 180mm. The coil removed for shimming. |
| Field integrals tuning done. |
| Long coil set back, PDI and XPS ON. |
| Undulator field integrals checked. If additional shimming needed, repeat coil measurements and tuning. |
| Field integrals shimming completed. |
| Field integrals measured for finding commissioning gap. |
| Coil and stages set in park position, PDI and XPS OFF, the coil removed. |
| Kugler bench ON and re-initialized. |
| MA position vs Y checked. |
| Trajectories and phase errors checked and within the tolerances. |
| Hall probe measurements to find commissioning gap done. (K=4.29) |
| Phase matching done. |
| Coil installed; optical measurements of coil position done. |
| Field integrals vs gap checked. |
| Final scan done, coil |
| Coil and stages set in park position. PDI and XPS OFF, the coil removed. |
| Gap clearance checked. Gap =7.350mm |
| Voltmeters checked. |
| Kugler bench ON and re-initialized. |
| Reference magnet measured. |
| MA position vs Y checked. Gap = 10mm |
| Half-gap encoders checked; Gap=10mm |
| Commissioning gap checked. |
| Final scan done, Hall probe. |
| PMs measured for fiducialization, file with results created. |
| Fiducialization done. |
| Touch probe measurements at 10mm gap done. |
| Touch probe measurements with taper +100µm, +300µm, -100µm done. |
| Capacitive sensor measurements, poles #10, 15, 101, 106 done. |
| Tuning Evaluation traveler created. |
| Vacuum chamber limit switches set. |
| Hard Stops set at 7.2mm Gap; Go / No-Go shim sizes recorded in a table below. |
| Controller Configuration saved. |

The following information is to be noted by an engineer:

| Average X (m): |  |
| --- | --- |
| Average Y (m): |  |
| Final Pitch (µrad): |  |
| Final Yaw (µrad): |  |
| Magnetic axis position X (m): |  |
| Magnetic axis position Y (m): |  |
| Coil X-position (m) |  |
| Coil Y-position (m) |  |
| Commissioning Gap (mm) |  |
| Reference magnet measurement (T) |  |
| Voltmeter X (1V) 1.00002 |  |
| Voltmeter Y (1V) 1.00002 |  |
| Voltmeter Y (6V) 6.00006 |  |
| Voltmeter Y (-6V) -6.00006 |  |
| US Half-Gap Encoder |  |
| DS Half-Gap Encoder |  |

Fiducialization:

The following information is to be noted by an engineer:

| Engineer (Initials): |  |
| --- | --- |
| Upstream RM X offset (m): |  |
| Upstream RM Y offset (m): |  |
| Downstream RM X offset (m): |  |
| Downstream RM Y offset (m): |  |
| Hard Stop US Go shim size |  |
| Hard Stop US No-Go shim size |  |
| Hard Stop DS Go shim size |  |
| Hard Stop DS No-Go shim size |  |

Have an alignment crew to measure offsets between RM and the undulator fiducials.

| Surveyor (initials): |  |
| --- | --- |
| Date (mm/dd/yyyy): |  |

Prepare the undulator for the move into the storage area:

- Power OFF the controller. Disconnect the network cable.

- Remove linear potentiometer brackets, if installed.

- Remove long coil brackets.

- Remove temperature sensors.

- Remove ground wire for capacitive sensors.

- Install all protective covers. Remove wooden platform between the undulator and granite.

- The undulator is ready for move into the storage area.

| Technician (initials) / date: |  |
| --- | --- |

URL for data:

V:\MET\MagServe\MagData\LCLS-II-HE\Undulator\HE-SXU_000\DATASET0003\Final Results\

| Undulator tuning completed (signed): |  |
| --- | --- |