## SLAC Traveler for LCLS Injector Dogleg Dipole Magnets

This traveler is intended to cover reception, preparation, mechanical fiducialization, and magnetic measurements of the two injector dogleg dipole magnets from Alpha Magnetics. These magnets are designated BX01 and BX02 and each has both main and trim coils.

Receiving:

The following information is to be noted upon receipt of the magnets by the SLAC MFD group:

| Received by |  | name |
| --- | --- | --- |
| Date received |  | mm/dd/yy |
| SLAC drawing number: | SA-380-300-00r0 |  |
| Vendor serial number from magnet label: |  | # |
| SLAC approved electrical safety covers? |  | Y/N |
| SLAC approved lifting eyes? |  | Y/N |
| Shipping Damage? |  | Y/N |
| Vendor tests passed on magnet label? |  | Y/N |

| Place barcode sticker on magnet and duplicate barcode sticker here: |  |
| --- | --- |

Preparation:

A beam direction arrow, with text “beam direction”, is to be applied to the top and connector side of the magnet, preferably by stenciling or rubber stamp, or by sticker supplied by LCLS.   If the magnet is viewed with connector box placed on the left, the beam direction is away from the viewer.

| Beam direction arrow applied |  | installer initials |
| --- | --- | --- |

Fiducialization:

Fiducialization may be done before or after magnetic measurements.

The magnet is to be fiducialized by the CMM group.  This will require the installation of removable tooling balls, location of the geometric axis of the poles of the magnet, and location of tooling balls with respect to the center of this geometric axis when the poles are aligned precisely horizontal.

| CMM technician |  | name |
| --- | --- | --- |

name of file and URL location of CMM data

Magnetic Measurements:

Upon receipt of the magnet, the following information should be recorded:

| Date of arrival to Mag. Meas. group: |  | mm/dd/yy |
| --- | --- | --- |
| Responsible measurement operator name: |  | name |

1)  Verify that the magnets are complete and undamaged, including wiring and water cooling connections.

| Incoming inspection OK |  | technician initials |
| --- | --- | --- |

Please measure the length-integrated field, (Bdl, over several current settings and the field harmonics as described below.  Please perform the following measurements:

- First determine the main-coil connection polarity, while main supply is outputting positive current, which produces a field polarity as shown below:

- Then mark the polarity near the main magnet leads with clear “+” and “” labels as shown above.

| Polarity labeled |  | technician initials |
| --- | --- | --- |

- Also mark the trim leads such that with the trim supply outputting positive current, the trim coil increases the magnetic field. The main and trim polarities are defined as “positive” for each of these two BX01 and BX02 magnets as described in PRD 1.1-010 (http://www-ssrl.slac.stanford.edu/lcls/prd/1.1-010-r0.pdf).

| Trim coils labeled |  | technician initials |
| --- | --- | --- |

- Set the magnet up for integral field measurements.  Connect the main magnet terminals (not the trim) to a power supply with maximum current I ( 340 A.   Leave the trim coil disconnected for now.

- Run the magnet up to 225 A for ~60 minutes to warm it up (record temperature).

| Ambient Temperature |  | degrees F |
| --- | --- | --- |
| Final Temperature |  | degrees F |

- Standardize the magnet, starting from 0 to 340 A and back to zero, through three full cycles, finally ending at 0, with a flat-top pause time (at 0 and 340 A) of 15 seconds.  If getting to zero requires turning the power supply off, do so in this and subsequent tests.

| Standardization complete |  | technician initials |
| --- | --- | --- |

- Maintaining this history cycle, measure (Bdl from 0 to 340 A in 15-A steps, including zero (23 ‘up’ measurements).  If the maximum integrated field is <2.0 kG-m, please record the current necessary to achieve this field and re-standardize up to the new current, starting the procedure again from that point.

| Current to get 2.0 kG-m |  | Amps |
| --- | --- | --- |

name of file and URL location of (Bdl data

- Still maintaining the history cycle, measure (Bdl back down from 340 A to 0 in 15-A steps, including zero (23 ‘down’ measurements).

name of file and URL location of (Bdl data

- With the main coils still hooked up, connect the trim coil to a bipolar 12-A supply with proper polarity as determined above.

- Still maintaining the history cycle, run the main coil up to 225 A, pause at least 15 seconds, and measure (Bdl as a function of trim coil current from 0 to +12 in 1-A steps, including zero (13 ‘up’ measurements), and again from +12 to 12 A in 1-A steps (25 ‘down’ measurements). Set the trim current back to 0.

name of file and URL location of (Bdl data

- Run the main coil directly up to 340 A, pause at least 15 seconds, and measure (Bdl as a function of trim coil current from 0 to +12 in 1-A steps, including zero (13 ‘up’ measurements), and again from +12 to 12 A 1-A steps (25 ‘down’ measurements).  Set the trim current to 0.

name of file and URL location of (Bdl data

- Set the main coil to 0 current by ramping down from 340 A at the same ramp rate used in the standardization cycle.  Again, if setting the coil to zero requires turning the supply off, do so.

- Pause at least 15 seconds at 0 A, and measure (Bdl as a function of trim coil current from 0 to 12 in 1-A steps, including zero (13 ‘down’ measurements), and again from 12 to 12 A in 1-A steps (25 ‘up’ measurements). Set the trim current to 0.

name of file and URL location of (Bdl data

- Set magnet up in rotating coil jig.  With the trim coil at zero current, measure field harmonics at 225 A, then again at 225 A, but now with the trim current at +12 A, and again at 340 A (with trim at 0 and again at +12 A).  Use a 1-inch diameter probe.

| Rotating coil designation |  | text |
| --- | --- | --- |
| Rotating coil radius |  | m |

names of file and URL location of field harmonic data

Upon completion of tests, send traveler to Paul Emma, at mailstop 103

15)  Approval and Assignment by Paul Emma:

| Magnet accepted: |  | Y/N |
| --- | --- | --- |
| Assigned Location |  | Mad Deck Name |

Upon completion, send traveler to Kathleen Ratcliff at mailstop 18.