## HXRSS Chicane Remnant Field
Date: June 2, 2020
Barcode | 1049 | 1048 | 1050 | 1051 | Notes: | By field meaurements made at z center, at the edge of the pole. The vacuum chamber fills the gap.
MAD Name | BXHS1 | BXHS2 | BXHS3 | BXHS4
Meas # | [5 empty] | Standardized then degauss |  |  | Measurements made by Scott Anderson and Kristi Luchini
1 | -0.4 | 0.8 | 0.6 | -0.9 |  | Run to + 6 Amps then degauss
2 | -0.6 | 0.9 | 0.7 | -0.3 |  | Run to -6 Amps then degauss
3 | -0.4 | 0.7 | 0.4 | -0.6 |  | Run to +6 Amps then degauss
4 | -0.5 | 0.9 | 0.4 | -0.4
Mean | -0.475 | 0.825 | 0.525 | -0.55 | 0.59375
Stdev | 0.095743 | 0.095743 | 0.15 | 0.26458
Date: July 27th 2021
Meas Location | BXHS1 | BXHS2 | BXHS3 | BXHS4
| (G) | (G) | (G) | (G) |  | The measurements on July 27th 2021 where done as is and in the below locations
1 | -1.7 | 1.8 | 2.3 | -1.7 |  | on the pole tip as far as the probe could go in, which only a mm or 2
2 | -1.7 | 2.2 | 1.6 | -1.7
3 | -1.8 | 2.6 | 2 | -1.7 |  |  |  | Core
4 | -1.6 | 2.5 | 1.8 | -1.8 |  | Upstream
5 | -1.8 | 1.1 | 1.5 | -1.7
Mean | -1.72 | 2.04 | 1.84 | -1.72 | 1.83 | 1 |  | Pole |  | 5 | Downstream
Stdev | 0.083666 | 0.61074 | 0.32094 | 0.044721 | 0.151
[7 empty] | 2 | 3 | 4
| 1.72 | 2.04 | 1.84 | 1.72 |  |  | Open Side of Magnet Gap
Date: November 8th 2023 | BXHS1 | BXHS2 | BXHS3 | BXHS4 |  | November 8th 2023 Measurements made by Scott Anderson at locations 2, 3 and 4
| (G) | (G) | (G) | (G)
Poe Tip Field Center of Magnet | -0.5 | 0.8 | 0.6 | -0.5
Error +/- | 0.1 | 0.1 | 0.1 | 0.1

## SXRSS Chicane Remnant Field
Date: July 2013 - Original Measurements
Meas Location | BXSS1 | BXSS2 | BXSS3 | BXSS4
| (G) | (G) | (G) | (G)
Poe Tip Field Center of Magnet | 0.6 | -0.5 | -0.4 | 0.5
Stdev | 0.02 | 0.02 | 0.02 | 0.02
Date: July 27th 2021 | [5 empty] | The measurements on July 27th 2021 where done as is and in the below locations
Meas Location | BXSS1 | BXSS2 | BXSS3 | BXSS4 |  | on the pole tip as far as the probe could go in, which only a mm or 2
| (G) | (G) | (G) | (G)
1 | 0 | -0.3 | -0.4 | 0.4 |  |  |  | Core
2 | 0 | -0.3 | -0.4 | 0.4 |  | Upstream |  |  |  | Downstream
3 | 0 | -0.3 | -0.4 | 0.4
Mean | 0 | -0.3 | -0.4 | 0.4 |  |  |  | Pole
Stdev | 0 | 0 | 6.7987e-17 | 6.7987e-17
[7 empty] | 1 | 2 | 3
Open Side of Magnet Gap
Date: November 8th 2023 | BXSS1 | BXSS2 | BXSS3 | BXSS4
| (G) | (G) | (G) | (G) | November 8th 2023 Measurements made by Scott Anderson at locations 2, 3 and 4
Poe Tip Field Center of Magnet | 0 | -0.3 | -0.4 | -0.1
Error +/- | 0.1 | 0.1 | 0.1 | 0.1

## HXR Chicane Degauss Hunt 9-21
Remnant Field |  | Last Current | # Points | f | Ramp Rate | Meas | Repeat | Repeat | Repeat | Repeat | Repeat | Comment | Data Taken 9/13/2021 -9/16/2021 by SDA
(G) |  | (A) |  |  | (A/s)
2.2 |  | -0.0114 | 52 | 0.88 | 2 | [6 empty] | Original
1.6 |  | 0.0075 | 51 | 0.87 | 2
2.2 |  | 0.0223 | 51 | 0.89 | 2
2.2 |  | 0.014 | 55 | 0.89 | 2
1 |  | 0.0057 | 53 | 0.87 | 2
1.8 |  | -0.0065 | 52 | 0.87 | 2
1.3 |  | -0.0199 | 52 | 0.89 | 2
3.2 |  | 0.013 | 51 | 0.88 | 2
Mean Remnant | StDev
0.066667 | 0.29439 | -0.0088 | 54 | 0.88 | 1 | -0.3 | 0.5 | 0.2 | 0 | -0.2 | 0.2
1.8 |  | -0.025089 | 50 | 0.89 | 2
0.05 | 0.49329 | -0.025089 | 50 | 0.89 | 1 | 0.3 | 0.6 | -0.2 | -0.5 |  |  | Last current is fairly large at -5.1 G
-2 |  | -0.0086 | 50 | 0.87 | 1
-1.5 |  | -0.0049 | 54 | 0.87 | 1
0.55 | 0.75939 | 0.0057 | 53 | 0.87 | 1 | -0.1 | 0.1 | 0.6 | 1.6
-0.275 | 0.73655 | -0.0114 | 52 | 0.88 | 1 | -0.6 | 0 | 0.6 | -1.1
-0.53333 | 0.73711 | -0.0114 | 55 | 0.88 | 1 | 0.3 | -1.1 | -0.8 |  |  |  | Start with mini Stdz -6 to 6 to -6
0.85 | 0.21213 | -0.025089 | 53 | 0.89 | 1 | 0.7 | 1 |  |  |  |  | Start with mini Stdz -6 to 6 to -6
-0.83333 | 1.5948 | 0.0057 | 56 | 0.87 | 1 | 1 | -1.6 | -1.9 |  |  |  | Start with mini Stdz -6 to 6 to -6
Note on H12: Anderson, Scott D.: Repeat Degauss straightway from previous value.
Note on I12: Anderson, Scott D.: After Stdardize and the ram to ~4 amps.
Note on J12: Anderson, Scott D.: After Stdardize and the ram to ~4 amps.
Note on K12: Anderson, Scott D.: After Standardize and ramping to ~4 A
Note on L12: Anderson, Scott D.: Repeat Degauss
Note on J17: Anderson, Scott D.: Repeat Degauss
Note on I19: Anderson, Scott D.: Repeat Degauss