# The questions about calibration

Chenwan

2011-12-13

- The calibration parameter:

The range of gaussmeter: 0.05÷2.1T

The range of Voltmeter ：10V

The style of voltmeter: HP3458A

The style of NMR: Metrolab PT2025 (5 probes)

The power supply stability: 3ppm

Temperature: 20.5

Calibration accuracy: 0.3G

- How to get the zero offset of the probe? Usually, we do not calculate and use probe zero offset. We start and end our Hall probe scans in zero gauss chambers and apply zero offset and zero drift corrections based on this measurements. You could find information about zero offset at the end of the .log file: “Estimated probe offset = “, but units will be V. You could convert it to T by dividing on the second polynomial coefficient.

- What’ the meaning of the titles?(Est.Field,voltage,field) .This is a raw data table. Since we have 5 NMR probes, we change them during the calibration. Data in this table are in the order they were measured. Sorted data are in the “summary” table, ready for plotting and calculations.

We have an extra Group 3 probe in the calibration magnet to roughly set the measurement range of the NMR. “Est. Field” is Group 3 probe field measurement. “Voltage” is HP3458A voltage measurement and “field” is corresponding NMR measurement.

Current :  the current of power supply.

T : the ambient temperature.

Est.Field: Group 3 probe measured field

Voltage: HP3458A voltage

Field: NMR field

| Current(A) | T(C) | Est.Field(T) | Voltage | Field |
| --- | --- | --- | --- | --- |
| 100 | 20.470291 | 1.7873 | 1.767095 | 1.771812 |
| 98 | 20.517058 | 1.77325 | 1.753133 | 1.757791 |
| 96 | 20.563316 | 1.75865 | 1.738526 | 1.743159 |

Summary

size 	 126

Voltage 	      Field

-1.771032 	 -1.772114

-1.757039 	 -1.758102