Meaurement Notes for LCLS Quad 002804

For stretched wire runs the PID values were set to the following:

P = 10, instead of the default 20

I = 50, instead of the default 40

D = 10, defualt 10.

This allowed smoother velocities for the stages.

Run 5.  Center run 5 shows a 0.0834 mm offset to the magnet center.  Since the run had no warmup I will not move the wire.

Run 6. Integrated Gradient at 6 amps after a 5 hour warm up.  GL =  3.9671 +/-  2.7920e-003 (std/mean = 7.0379e-004) using Matlab code Plot_Stretchedwire_BLvsX_Data which corrects for the current before fitting.   Current = / 6.005210 Amps  +/- 2.10600e-005

GL/I =  3.9671 T / 6.00521025 Amps =  0.66061 +/-  4.64936e-4 T/ A

Run 8. Measurement of the Hcor Int Field Strength BL/I =  6.7791e-004   +/-  4.7010e-006  (std/mean =  6.9346e-003)  at  2.0001   Amps

Run 15.  Trip test at 6 amps with 3 stdz cycles.  Integrated Gradient was measured twice, then the magnet was tripped and the magnet was stdz 3 times and measured twice.  Results SL/I before = 0.66405945 Tm/A, SL/I after trip and resdtz = 0.664502.   SL/I After/Before =  1.0006664

Run 16.  Long Integrated Gradient run at 6 amps.  Mean of last 10 measurements with the mirror plates on mean_SL/I = 0.66137672.

Using Old Data runs 22 (last 15 measurements) and run 26 (and accounting for different coil constants) the Int Grad over I with no mirror plates is is  0.664289.  The ratio to 0.664289/ 0.66061 = 1.0055706.   SL with No Mirror at 6 Amps = 0.664289 T/A *  6.00 A =  3.9856794 +/- 0.00280511 T.

Run 17, 18 & 19 trying to determine how much the center shifts when start cold vs restarted after a few minutes Last Ctr measurement

Run 18.

First Ctr Measurement after 36 minutes off. Last Ctr measurement when it stopped Sunday morning.

Run 19.  First Measurement on Monday

| Run | Date/Time | Imag | Xctr | Yctr | BL Coil | TR Coil | Mag Steel | Ambient |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 17 end | Friday 11:26:20 | 6.004 | -248.79 | 210.22 | 32.00 | 32.44 | 28.23 | 21.66 |
| 18 1st | Friday 12:04:24 | 6.005 | -248.56 | 207.86 | 27.24 | 27.76 | 27.06 | 21.40 |
| 18 end | Sunday 06:23:54 | 6.005 | -246.06 | 211.08 | 31.97 | 32.59 | 28.12 | 21.48 |
| 19 1st | Monday 13:02:02 | 6.005 | -241.29 | 190.92 | 22.56 | 22.37 | 21.34 | 21.47 |
| 19 2nd | Monday 13:08:24 | 6.005 | -241.42 | 191.46 | 24.37 | 24.05 | 21.52 | 21.45 |
| 19 End | Monday 15:47:37 | 6.005 | -241.75 | 200.55 | 30.87 | 31.10 | 27.04 | 21.44 |

Run 20 Was the before magnet split, run was overnight and was terminated right before split was done.  The Heidenhain gauge and a digital dial gage were put magnet base and two other digital gages were put on the magnet itself.  Before the measurement, all gauges were zeroed.

Run 21 was done after split.   The magnet moved horizontally -8 microns with respect to the coil and the base moved +1 micron with respect to the coil, so the magnet moved -7 microns horizontally with respect to the coil during the split.  The vertical movement was not measured though it should be able to change because it is sitting on a rigid base.

| Run | Date/Time | Imag | Xctr | sigXctr | Yctr | Error Yctr |
| --- | --- | --- | --- | --- | --- | --- |
| 20 end | Tuesday 13:22 | 6.004 | -240.72 | 0.08 | 204.73 | 0.08 |
| 21 1st | Tuesday 12:04:24 | 6.005 | -246.45 | 0.07 | 206.16 | 0.05 |
| Gage offset |  |  | -7 | 0.5 | N/A | N/A |
| Difference before split to after split |  |  | 1.27 | 0.75 | -1.43 | 0.13 |

Run 22.  Magnet moved to magnetic center using data form run 21.   Heidenhain read 185.5 microns and Mititoyo gage read 141 microns.  Avg = 163.  Coil measured the move as 159 microns (-246.70 start position, -87.74 end position).

Run 23 was measuring Integrated Gradient with the mirror plates on to check the Coil Constant (1174.9497).  Result was 0.66061 (wire) / 0.660476845 (coil)  = 1.000201603  which is within error of wire.

Run 28 measured the Integrated Gradient from -6 to 6 to -6 Amps after a 2 hours warmup and then after 5 hour warmup.  The mean of the ratio of the 5 hour run to the 2 hour run for all non zero currents is 1.00004119.

Found reference radius, for the main coil, to be in correct.  The radius used for all the harmonics measurements was 0.004521 m but it should have been 0.004114984 m.  The correct radius was calculated using run 24 and run 25 strdat data at -6 amps.  The main coil CC = 2461.286927, Int Str at -6 amps using main coil = 0.6871252.    Int Str at -6 amps using double (calibrated) coil = 0.6594735, so the new CC for the main coil is 2461.286927*0.6594735/ 0.6871252 = 2362.238358.  The radius from this CC is 0.004114984 m, assuming 50 turns.  To calculate the change in harmonics I used Matlab function  Harmonics_Radius_Change_old and used the 0.004114984 as the old and ,0.004521 as the new.  I also checked these numbers by using the voltages from the logfile. Harmonics_Radius_Change_old(‘hardat25 +6 amps.txt’, 0.004114984,0.004521, 2) was the command.

This changes the harmonics (makes them larger) .  As an example here are the harmonics with the 0.004521 radius at 6 amps for run 25 (the last measurement):

n        BLn         sigBLn     THspole   sigTH    BLn/BLN   sBLn/BLN

(Tm)          (Tm)       (deg)    (deg)      (%)       (%)

---  ------------+------------  --------+--------  ---------+---------

1    3.5381e-04   8.2410e-08     52.72     0.01    2.38306   0.00056

2    1.4847e-02   5.3242e-07     36.97     0.00  100.00000   0.00507

3    1.0586e-04   1.8938e-07     39.19     0.05    0.71299   0.00128

4    2.5106e-05   2.8630e-07    -44.38     0.07    0.16910   0.00193

5    5.9268e-06   1.9457e-07     33.67     0.35    0.03992   0.00131

6    3.8810e-05   1.5325e-07    -25.33     0.05    0.26140   0.00103

7    6.7484e-06   2.4388e-07     24.30     0.18    0.04545   0.00164

8    1.0852e-05   2.6056e-07     19.39     0.10    0.07309   0.00175

9    5.5970e-06   1.7397e-07     16.27     0.21    0.03770   0.00117

10    1.6569e-05   3.1007e-07     16.33     0.15    0.11160   0.00209

11    3.6714e-06   2.0339e-07     11.76     0.29    0.02473   0.00137

12    1.4869e-05   4.5868e-07     10.57     0.15    0.10015   0.00309

13    1.1641e-05   3.0627e-07    -12.26     0.16    0.07841   0.00206

14    1.5453e-05   5.4646e-07      8.50     0.12    0.10408   0.00368

15    9.8928e-06   5.1299e-07     -9.86     0.14    0.06663   0.00345

16    2.3214e-05   5.7650e-07      7.84     0.09    0.15636   0.00388

Here are the correct harmonics at 0.004114984 using program:

Harmonics_Calc_from_VT('Run 25 +6 amp blhar from log.txt', 0.004521, 0.004114984,0.004521, 0.004114984,2)

Magnet Current, Imag = 6.004808 +- 0.000053 A

Harmonics given at R = 0.004115 m

Coil Radius, Rcoil = 0.004115 m

Main Harmonic N = 2

n        BLn        sigBLn     THspole   sigTH     BLn/BLN   sBLn/BLN

(Tm)         (Tm)       (deg)    (deg)       (%)       (%)

---  ------------+------------  --------+--------   --------+--------

1   3.8870e-004  1.0987e-007     52.72     0.01    2.38294   0.00068

2   1.6312e-002  5.9458e-007     36.97     0.00  100.00000   0.00515

3   1.1627e-004  2.4743e-007     39.19     0.05    0.71280   0.00152

4   2.7603e-005  3.2932e-007    -44.38     0.07    0.16922   0.00202

5   6.5072e-006  2.1973e-007     33.67     0.35    0.03989   0.00135

6   4.2661e-005  1.4151e-007    -25.33     0.05    0.26153   0.00087

7   7.3701e-006  2.1708e-007     24.30     0.18    0.04518   0.00133

8   1.1894e-005  3.0200e-007     19.39     0.10    0.07292   0.00185

9   6.1858e-006  2.1973e-007     16.27     0.21    0.03792   0.00134

10   1.8247e-005  3.4014e-007     16.33     0.15    0.11187   0.00208

11   3.9811e-006  2.1973e-007     11.76     0.29    0.02441   0.00134

12   1.6326e-005  5.1264e-007     10.57     0.15    0.10009   0.00314

13   1.2795e-005  3.6411e-007    -12.26     0.16    0.07844   0.00223

14   1.6967e-005  6.1450e-007      8.50     0.12    0.10402   0.00377

15   1.0912e-005  5.6970e-007     -9.86     0.14    0.06690   0.00349

16   2.5518e-005  6.2425e-007      7.84     0.09    0.15644   0.00383

Calculating the Int Str for the quad from the harmonics = 3.9640494 T using command:

Int_Str = Int_Str_from_Harmonic(2, 1.6312e-002, 0.004114984)