#### 1/19/2016

B9V2 Measurement at SLAC – Rev. 0

Domenico Dell’Orco

- Conditioning Cycle

Before starting any magnetic measurement remove the stiffeners at both ends of the insert that cover the V slot.

Raise the current from 0 to 265 A (230 A design current + 15%) in one minute, keep it at 265 A for two minutes and return the current to 0 in one minute. Repeat this cycle for 3 times. Make sure to follow the polarity marks. The B field is horizontal and directed from the V slot toward the magnet gap.

- Transfer function on the injected beam – Translating wire

Position the translating wire inside the pole aperture at 3 mm away from the INSERT flat surface near the V slot. The wire should be positioned at y= +63.45 mm at beam entrance and y=-8.20 mm (NEGATIVE SIGN) at the beam exit at z=+/- 576.83 mm (flush with pole, NOT flush with magnet insert). Y=0 is at the vertical center of chamber. See attachment. Move the translating wire by +/-1 mm vertically during measurements.

Measure the integrated field (Integral(Bxdz) vs. Ic while ramping the current from 0 A to 255 A (230 A design current + ~11%) and back to 0 A at increments of 15 A. Determine by interpolation the nominal current that provides an integrated field of 1.5341 Tm. The nominal current should be close to the 230 A design current.

- Leakage field measurement on the stored beam – Translating wire

The leakage field is mainly horizontal. The magnet is fully symmetric and no considerable vertical field should exist.

The stored beam is 15 mm from the V, the insert is 2.5 mm thick at the V, the injected beam is 3 mm from the pole. Total distance between Store Beam and Injected Beam is 20.5 mm (15+2.5+3).

Position the wire in the center of the V slot (Y=0, X= 17.5 mm from pole surface) and measure the integrated leakage at the following points:

X (from the pole surface): from 42.5 mm (25+15+2.5) to 12.5 mm in steps of 2.5 mm (or 3.0 mm if possible without hitting the V) and from 12.5 mm to 3 mm in steps of 0.5 mm. 32 data points. Wire position within +/- 0.2 mm in x. Wire vertical movement at x= 3 mm (0.5 mm from bottom of V) could problematic.

Measure the leakage field varying the current from 0 A to 180 A in steps of 20 A and from 180 A to 255 A in steps of 5 A (24 data points).

The expected integrated field leakage is ~15 Gm at 230 A at x=17.5 mm (stored beam).

Magnet Parameters (the measured insert length is 46.910”)

| 5/15/2015 -SPEAR3 SEPTUM - DD | B9V2 Final Design Review |
| --- | --- |
| Nominal Energy (GeV) | 3.0000 |
| Angle (deg) | 8.7900 |
| Magnetic arc length (m) | 1.1672 |
| Bending radius (m) | 7.6080 |
| Field (T) | 1.3144 |
| Magnetic straight length (m) | 1.1626 |
| Integrated field (T*m) | 1.5341 |
| ORBIT & GAP |  |
| Entrance height (m) | 0.0894 |
| Core entrance-vertex (m) | 0.5734 |
| Vertex-core exit (m) | 0.5802 |
| Core straight Length (m) | 1.1536 |
| Pole extender length (m) (each) | 0.0000 |
| YOKE |  |
| Core Height (m) | 0.7366 |
| Core Width (m) | 0.3362 |
| Core Weight (kg) | 2143 |
| Pole width (m) | 0.2540 |
| Gap Height (m) | 0.0090 |
| CONDUCTOR (nominal size) |  |
| Conductor width (mm) | 6.50 |
| Conductor height (mm) | 6.50 |
| Cooling Hole Diam (mm) | 4.50 |
| Conductor fillet radius (mm) | 1.00 |
| Cond. Cross-Sect. Area (mm2) | 25.49 |
| COIL (copper) | 1 coil |
| Turns per Pole | 42 |
| Pancakes per Pole | 1 |
| Conductor length/pole (m) | 137.7 |
| Total Coils Weight (kg) | 30.9 |
| MAGNET |  |
| Magnet efficiency (%) | 97.5 |
| Amp Turns per Pole | 9655.1 |
| Current (A) | 229.88 |
| Current density (A/mm2) | 9.02 |
| Inductance (H) (MAX) | 0.071 |
| Mag. Res. @ 40 C (mOhm) | 100.4 |
| Power/Magnet (kW) | 5.3 |
| Voltage drop/Magnet (V) | 23.1 |
| WATER FLOW at 21C |  |
| Water Circuits | 3 |
| Water Flow Rate/Magnet (gpm) | 2.0 |
| Water Pressure Drop (psi) | 150 |
| Water Temperature Rise (C) | 10.2 |
| Water speed (m/s) | 2.6 |
| Reynold's number | 11931 |

# Text boxes

1 min.

2 min

1 min.

0

0

265A