Magnet Measurement Test Plan

Fill

Logistic Info

| Person Responsible | Domenico Dell’Orco |
| --- | --- |
| Phone # | 650-926-8683 (cell. 510-219-0378) |
| Start Date | 02-24-2006 |
| Finish Date | 04-17-2006 |
| Acct # for alignment | 94-4306-1 |

Magnet Designation and Type

| Project Name | SSRL BL5 Skew quad magnet |
| --- | --- |
| Magnet Type | Skew Quad |
| Magnet Name/Number | BL5 |
| Serial # / PC # |  |
| Trim Coils | No | Trim Types |  |
| Magnet gap or diameter (mm) | 2.45 inches |

Alignment

| Coarse (1.0 mm)/Precision (0.05 mm) | Precision |
| --- | --- |
| Has the magnet been fiducialized | Yes |
| If it has not been fiducialized, what is origin point of the magnet, what is its coordinate system and what surfaces should be used for alignment? |  |
| + Z direction of magnet (Please provide drawing) |  |

Special Alignment Tolerances

| X |  |
| --- | --- |
| Y |  |
| Pitch |  |
| Yaw |  |
| Roll |  |

Water and Power.

| Water Pressure | 100-150 psi DP (1 water circuit) | PSI |
| --- | --- | --- |
| Water Flow | 0.5 GPM @ 150 psi DP | Gpm |
| Maximum Current Main | 30 A | Amps |
| Max Current Trims | N/A | Amps |

Polarity

| Main Coil Polarity. (Nearest North pole to Horizontal look in +z direction) | Power leads are marked. Positive polarity: B field coming out of the two coils vertically points toward the horizontal plane. |
| --- | --- |
| 1st Trim Polarity (Nearest North pole to Horizontal look in +z direction) | N/A? |
| 2nd Trim Polarity (Nearest North pole to Horizontal look in +z direction) | N/A |

Temperature Tests

| Temperature and Resistance Test | No |
| --- | --- |
| Areas to be measured | Check coil temperature far from cooling plates at maximum current. |
| Currents (Amps) | 30 A |
| Check water flow | Yes |

Standardization

| Magnet warm up. ? time (mins) @ ? current (Amps) | NONE. Iron free magnet. |
| --- | --- |
| Conditioning Standardization Cycles Performed before measurements: | None |
| Number of Stdz cycles | 0 |
| Stdz Ramp Rate (A/s) |  | Minimum Current (A) |  |
| Stdz Settle Time (sec) |  | Maximum Current (A) |  |
| Ramp Type (Linear or Three Linear or Cosine | Three Linear |

| Measurement Ramp |
| --- |
| Ramp Rate (A/s) | 3 | Settle Time (sec) | 5 |

| Integrated Strength | Main | Yes | Trims | N/A |
| --- | --- | --- | --- | --- |
| Currents: | 3, 6, 9, 12, 14.23, 15, 18, 21,24, 27, 30 A |  |
| Harmonics | Main | Yes | Trims | N/A |
| Currents: | 3, 6, 9, 12, 14.23, 15, 18, 21,24, 27, 30 A |  |

| Harmonic Tolerances (as % of main, Main =100 %) |
| --- |
| Dipole n = 1 | Quadrupole n = 2 | Sextupole n = 3 | Octupole n =4 |
| N/A | 1300*x G*cm (at 14.23 A) expected quad component of integrated field | No sextupole component expected if magnet is centered | 12.987*x^3 G*cm (at 14.23 A) expected octupole component of integrated field |

Field Map

| Type of map? (Transverse, axial, 3D, Locate Center) | Locate Center 1) Set the magnet a +45 degree, set current to 14.23 A, scan horizontally to locate the point with zero integrated field. Have the alignment crew measure location of vertical plane @ zero integrated field respect to fiducials. 2) Set the magnet a -45 degree, set current to 14.23 A, scan horizontally to locate the point with zero integrated field. Have the alignment crew measure location of vertical plane @ zero integrated field respect to fiducials. The intersection of the 2 planes defines the quadrupole center. |
| --- | --- |
| Dimensions ( X, Y, Z) | -25 mm < x < 25 mm |
| Spacing between step (X, Y, Z) | Dx = 1mm |

Notes and Additional Tests:

- Set the magnet at +22.5 degrees and measure integrated field from 0 A to 30 A (3A increments plus the 14.23 A nominal current) within a radius of 25 mm.