Magnetic test of the elliptically polarizing undulator for SPEAR BL05

#### Final testing

#### June 2012

##### Danfysik A/S

Gregersensvej 8

2630 Taastrup

Denmark

# Table of Contents

Table of Contents	2

Introduction	3

A.	Reproducibility of Hall probe bench	4

B.	Reproducibility of the integrating coil	7

C.	Magnetic alignment using the Hall probe and field roll-off	10

Magnetic Measurements without tunecorrectors	12

D.	Hall probe scans on the undulator axis	12

E.	Summarizing table of Hall results	19

F.	The first integrals	21

G.	First derivative of the first integrals	38

H.	The second integrals	49

Magnetic Measurements with tunecorrectors	59

I.	First integral scans after tune shim installation	60

J.	Hall probe scans on the undulator axis(13 mm gap)	65

K.	Analysis of peak field rolloff	67

L.	Reproducibility of the magnetic array	68

M.	Correction coils HP mode	72

General Comments	74

N.	Mounting of the permanent magnets and the shimming process	74

O.	Conclusion	77

## Introduction

We report the results of the magnetic measurements carried out at Danfysik of the 2m EPU for the SPEAR BL05 at SSRL.  The measurements were carried out using the Danfysik A/S Hall probe bench and flip coil bench. The resolution of the measurements at 13 mm gap is investigated in Section A of this report. The parameters for the measurements are given below:

Length of Hall probe scans				3300 mm

Distance between measuring points			1 mm

Length of flip coil					3500 mm

Diameter of flip coil					5.5 mm

Number of turns in flip coil				20

Interval measured by the flip coil			|x| ≤ 25 mm

Distance between flip coil measurement points		1 mm

The (x, y, z) coordinate system used in this report has the z-direction in the beam direction. In the raw data files it is, however, labeled as (x, z, s) where z is the vertical and s the longitudinal component.

The Hall probe scans were analyzed using the B2E program from ESRF based on the Igor Pro software. An additional set of macros has been written at Danfysik which makes use of the B2E program for the data analysis and then produces the result graphs. Nearly all the graphs presented in this report are made with these macros. The Hall probes were calibrated using a calibration magnet with a NMR probe as reference, as described in a separate calibration report. The calibration range is from –1.8 T to 1.8 T.  For this device, we had to design a special probe, which allowed for the calibration of the horizontal probe, up to 1 T.

From the Hall probe measurements we get:

Period length						139.98 mm ( 0.01 mm)

Total number of poles(including endpoles)				31

Measurements were carried out for gap sizes of 13,20,30,40, and 120 mm

In Section A and B of this report the resolution of the measurement equipment is investigated. Section C concerns the alignment of the undulator and the measurement equipment. Section D contains a summary of the on-axis measurements with the Hall probe, as obtained for the device without tune correctors installed.   The integrated multipoles and the second integrals as a function of the x-position was measured with the flip coil and the results of the analysis can be found in Sections F,G, and H.  In sections I and J, the magnetic results for the device after tuneshim installation are shown. We have performed some mechanical hysteresis tests, and the results of these are shown in section L.

General comments and information on the mounting of the permanent magnets on the undulator and the shimming procedure is given in Section N.  Finally a conclusion is given in Section O.

## Reproducibility of Hall probe bench

The reproducibility of the measurement equipment is investigated using the five on-axis Hall probe measurements in the HP Mode and the flip coil measurements at 13 mm gap.

The 13 mm gap was chosen to check the reproducibility of the Hall probe bench as the gradients are largest at the minimum gap and the effects of differences between the real and recorded Hall probe position will be maximum. Such differences might occur due to the resolution in the Heidenhain ruler used to read the Hall probe position, vibration in the timing belt used to move the probe and non-linearity in the rails used to guide the probe. Five Hall probe scans have been done under identical conditions and analyzed individually. In the Figure 1 to Figure 5, each of the scans is identified with a separate color.

Figure 1 shows the vertical angle for a 3 GeV electron. The reproducibility is good due to the small value of the transverse field on the undulator axis. The reproducibility in the horizontal angle averaged over one period is shown in Figure 2. The fine structure is well reproduced from scan to scan. The values agree at the start and end of the scans because of the normalization to the flip coil values for the first integral, and in the middle of the undulator there is only up to about 2 radian difference between the highest and lowest scan.

The orbits are shown in Figure 3 and Figure 4. Both the vertical and horizontal orbit reproduces well and the spread in the orbit walk is around than 0.5 m  for the vertical orbit and 1.3 m  for the horizontal orbit. The phase angle error at each pole is shown in Figure 5.

Figure 1 The vertical electron angle as obtained from 5 hall probe scans measured at 13 mm gap, in the HP mode for x=0 and y=0.

Figure 2 The horizontal electron angle, averaged over on undulator period, as obtained from 5 hall probe scans measured at 13 mm gap, in the HP mode, for x=0 and y=0.

Figure 3 Vertical orbit walk from 5 hall probe scans measured at 13 mm gap, in the HP mode, for x=0 and y=0.

Figure 4 . Horizontal orbit walk from 5 hall probe scans measured at 13 mm gap, in the HP mode, for x=0 and y=0.

Figure 5 Phase angle error calculated from 5 hall probe scans measured at 13 mm gap, in the HP mode for x=0 and y=0.

## Reproducibility of the integrating coil The reproducibility of the flip coil has been determined from five normal and five twisted coil scans measured at 13 mm gap ### First Integral Figure 6 Reproducibility of the horizontal 1st field integral, from the flip coil Figure 7 Reproducibility of the vertica 1st field integral, from the flip coil ### Second integral Figure 8 Reproducibility of the horizontal 2nd field integral, from the flip coil. Figure 9 Reproducibility of the vertical 2nd field integral, from the flip coil.

## Magnetic alignment using the Hall probe and field roll-off

The vertical position of the median plane was measured for each horizontal pole at 50  mm gap by making Hall probe scans at y = -10, -5, 0, 10 and 5 mm, with the device placed in the CP mode. The horizontal position was determined from Hall scans at x = -20, -10, 0, 10, 20 mm. The scans were analyzed using an Igor Pro macro, which fits a parabola to the measured peak field values for each pole and determines the y (or x) position of the minimum and the fit error. The data are plotted as a function of the pole number and a line is fitted to the data giving the position of the mid plane. The first and last two poles are excluded from the fit. The parameters of the line fit:

y = A + B*(Pole #)   or     x = A + B*(Pole #)

are given in Table B1 and B2. The scattering of the magnetic center positions is primarily due to the internal inhomogeneity of the magnetic blocks and cannot be avoided, unless actions are taken to physically correct this by transverse shift of the magnets. The observed scattering is very reproducible and the uncertainties on the linear fit parameters are relatively small. The uncertainty on each pole center as determined from the polynomial fit is indicated in the graphs by the error bars (not including uncertainty on the measurements).  The parameters for the horizontal and vertical position of the undulator mid-plane are given in Tables B1 and B2. These results show that we have a good alignment of the Hall probe is traveling with respect to the optimal magnetic center axis of the device.

Table B1. Parameters for the horizontal position of the undulator mid plane.

| Gap (mm) | A (mm) | B (mm/pole #) | Center off-set (mm) |
| --- | --- | --- | --- |
| 50 | -0.08 ± 0.1 | 0.00 ± 0.00 | -0.0312 |

Table B2. Parameters for the vertical position of the undulator mid plane.

| Gap (mm) | A (mm) | B (mm/pole #) | Center off-set (mm) |
| --- | --- | --- | --- |
| 50 | 0.01± 0.06 | 0.003 ± 0.00 | -0.03 |

Figure 10 Horizontal alignment analysis

Figure 11 Vertical alignment analysis

# Magnetic Measurements without tunecorrectors

The first part of the measurement programme, concerns the magnetic measurements of the device without tuneshims.

## Hall probe scans on the undulator axis

### HP Mode: 13 mm gap

Two hall probe scans along the center axis of the device have been measured at 13 mm gap, in each of the 4 modes.  The data is analyzed with B2E in Igor Pro using a macro made by Danfysik. The Igor Pro files are saved together with the data files.

From the VP hall probe scans we obtain an average period length of u = 139.98±0.001mm from the position of the 1. harmonic in the Fourier transformed of the horizontal field data.

The following figures  show the field components, electron angle and orbits, vertical peak field of the poles and phase error at each pole, obtained from a Hall probe scan at 13 mm gap on the center line.

Figure 12 The By-component of the magnetic field at 13 mm gap for x = y = 0(HP Mode).

Figure 13 The Bx-component of the magnetic field at 16 mm gap for x = y = 0(HP Mode).

Figure 14 The electron angles, averaged over on undulator period, at 13 mm gap for x = y = 0(HP Mode).

Figure 15 The electron orbits for a 3 GeV electron at 13 mm gap for x = y = 0(HP Mode).. The horizontal orbit averaged over one undulator period is also shown.

Figure 16 The phase error from vertical field as a function of pole number for x=0 and y=0(HP Mode), at 13 mm gap

Figure 17 The calculated flux spectra compared with the case of a perfect planar undulator without phase error.(HP Mode)

### CP+ Mode: 13 mm gap

Figure 18 The Bx- and By-component of the magnetic field at 13 mm gap, for x=y=0(CP+ mode)

Figure 19 The electron orbits for a 3 GeV electron at 13 mm gap for x = y = 0(CP+ Mode).. The horizontal orbit averaged over one undulator period is also shown.

### CP- Mode: 13 mm gap

Figure 20 The Bx- and By-component of the magnetic field at 13 mm gap, for x=y=0(CP- mode)

Figure 21: The electron orbits for a 3 GeV electron at 13 mm gap for x = y = 0(CP- Mode).. The horizontal orbit averaged over one undulator period is also shown.

### VP Mode( 30 mm gap)

Figure 22 The Bx-component of the magnetic field at 16 mm gap for x = y = 0(VP Mode).

Figure 23 The electron orbits for a 3 GeV electron at 13 mm gap for x = y = 0(VP Mode).. The horizontal orbit averaged over one undulator period is also shown.

## Summarizing table of Hall results

Table 1 Gap summary for the HP mode

| On axis Gap | Ix | Iy | Jx | Jy | Bz_peak | RMS(Bz) | Ephoton |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  | Gm | Gm | Gm^2 | Gm^2 | T | Degrees | eV |
| 13 | -0.4 | 0.28 | 0.1374 | -0.4028 | 0.9015 | 2.95 | 7.662 |
| 20 | -0.26 | 0.26 | 0.3634 | -0.4481 | 0.7917 | 2.77 | 10.341 |
| 30 | -0.17 | 0.22 | 0.4234 | -0.3461 | 0.6436 | 2.71 | 16.231 |
| 40 | -0.1 | 0.2 | 0.4951 | -0.2781 | 0.5129 | 2.75 | 25.664 |
| 120 | 0 | 0 | 0.4544 | -0.2220 | 0.0727 | 2.79 | 417.918 |

Table 2 Gap summary for CP+ mode

| On axis Gap | Ix | Iy | Jx | Jy | Bz_peak | Bx_peak | Ephoton |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  | Gm | Gm | Gm^2 | Gm^2 | T | T | eV |
| 13 | 0.02 | 1.04 | 1.6457 | 1.3214 | 0.7335 | 0.7317 | 7.401 |
| 20 | -0.04 | 0.86 | 0.8960 | 0.8425 | 0.5553 | 0.5549 | 11.970 |
| 30 | -0.04 | 0.74 | 0.5360 | 0.5636 | 0.3886 | 0.3866 | 23.163 |
| 40 | -0.08 | 0.73 | 0.2816 | 0.7201 | 0.2754 | 0.2745 | 43.915 |
| 120 | 0.01 | 0.11 | 0.4485 | 0.0341 | 0.0209 | 0.0211 | 564.026 |

Table 3 Gap summary for CP- mode

| On axis Gap | Ix | Iy | Jx | Jy | Bz_peak | Bx_peak | Ephoton |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  | Gm | Gm | Gm^2 | Gm^2 | T | T | eV |
| 13 | -0.24 | 0.51 | 0.0641 | 0.1678 | 0.7321 | 0.7335 | 7.395 |
| 20 | -0.04 | 0.57 | 0.9012 | 0.1090 | 0.5539 | 0.5557 | 11.971 |
| 30 | 0.03 | 0.6 | 1.2141 | 0.2200 | 0.3878 | 0.3872 | 23.163 |
| 40 | 0 | 0.51 | 1.0361 | 0.2367 | 0.2750 | 0.2748 | 43.932 |
| 120 | -0.03 | 0.12 | 0.5176 | 0.1262 | 0.0209 | 0.0211 | 563.957 |

Table 4 Gap summary for VP mode

| On axis Gap | Ix | Iy | Jx | Jy | Bx_peak | RMS(Bx) | Ephoton |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  | Gm | Gm | Gm^2 | Gm^2 | T | Degrees | eV |
| 30 | -0.03 | 0.38 | 0.9492 | 0.0192 | 0.4744 | 4.08 | 30.067 |

## The first integrals

The normal and skew first integrals have been measured in the interval –2.5  x  2.5 cm in 0.1 cm steps with the flip coil at the specified phases at 13 mm gap, and at increasing gaps. Using an Igor Pro macro the averages of the measured field integrals were calculated.   The data were then converted from Gm to Gcm and the x-positions converted from mm to cm.   Multipole analysis was carried out for this device, and limits were imposed on the multipole content within a  25 mm window.  As we can see from Table 5 to Table 8, the multipole content is outside these limits at some of the chosen phases, only at minimum gap.

In the following figures, the first integrals are presented for the 4 phases at 13 mm gap, as well as the first integrals for the HP mode at increasing gaps.  We see that, as the magnetic gap is increased, the integrals decay rapidly.  Included in the graphs are the polynomial fits, which which we can deduce the multipole content.  At minimum gap, the multipoles are just outside the specification, due to a rapid field integral variation near the undulator axis.

Figure 24 First integrals. 13 mm gap. HP mode

Figure 25 First integrals. 20mm gap. HP mode

Figure 26: First integrals. 30mm gap. HP mode

Figure 27 First integrals. 40mm gap. HP mode

Figure 28 First integrals.120mm gap. HP mode

Table 5: Multipoles at  increasing gaps in HP mode.

| Gap (mm) | Mode | Multipole | Skew | Skew error | Normal | Normal error |
| --- | --- | --- | --- | --- | --- | --- |
| 13 | HP | 2-pole (Gcm) | -33.9 | 2.4 | 27.9 | 2.4 |
|  |  | 4-pole (G) | 65.5 | 5.4 | -78.4 | 5.6 |
|  |  | 6-pole (G/cm) | 10.7 | 10.8 | 93.3 | 11.1 |
|  |  | 8-pole (G/cm^2) | -50.4 | 9.5 | 67.3 | 9.8 |
|  |  | 10-pole (G/cm^3) | 5.4 | 12.5 | -59.5 | 12.9 |
|  |  | 12-pole(G/cm^4) | 14.6 | 5.3 | -26.3 | 5.5 |
| Gap (mm) | Mode | Multipole | Skew | Skew error | Normal | Normal error |
| 20 | HP | 2-pole (Gcm) | -28.4 | 0.7 | 39.7 | 1.3 |
|  |  | 4-pole (G) | 26.4 | 1.1 | -21.9 | 2.0 |
|  |  | 6-pole (G/cm) | 14.0 | 0.7 | 4.7 | 1.2 |
|  |  | 8-pole (G/cm^2) | -7.5 | 0.7 | 7.6 | 1.2 |
|  |  | 10-pole (G/cm^3) | -0.6 | 0.1 | -0.8 | 0.2 |
| Gap (mm) | Mode | Multipole | Skew | Skew error | Normal | Normal error |
| 30 | HP | 2-pole (Gcm) | -20.5 | 0.3 | 29.7 | 0.4 |
|  |  | 4-pole (G) | 12.6 | 0.3 | -4.6 | 0.3 |
|  |  | 6-pole (G/cm) | 10.2 | 0.3 | 2.1 | 0.4 |
|  |  | 8-pole (G/cm^2) | -2.0 | 0.1 | -0.8 | 0.1 |
|  |  | 10-pole (G/cm^3) | -0.6 | 0.1 | -0.4 | 0.1 |
| Gap (mm) | Mode | Multipole | Skew | Skew error | Normal | Normal error |
| 40 | HP | 2-pole (Gcm) | -9.4 | 0.2 | 24.4 | 0.1 |
|  |  | 4-pole (G) | 6.6 | 0.2 | -4.3 | 0.1 |
|  |  | 6-pole (G/cm) | 6.4 | 0.2 | 1.3 | 0.1 |
|  |  | 8-pole (G/cm^2) | -1.8 | 0.0 | -0.6 | 0.0 |
|  |  | 10-pole (G/cm^3) | -0.4 | 0.0 | -0.2 | 0.0 |
| Gap (mm) | Mode | Multipole | Skew | Skew error | Normal | Normal error |
| 120 | HP | 2-pole (Gcm) | 0.0 | 0.0 | 0.0 | 0.0 |
| Gap (mm) | Mode | 4-pole (G) | -0.0 | 0.0 | -0.0 | 0.0 |
| 120 | HP | 6-pole (G/cm) | 0.0 | 0.0 | -0.0 | 0.0 |
|  |  | 8-pole (G/cm^2) | 0.0 | 0.0 | 0.0 | 0.0 |
|  |  | 10-pole (G/cm^3) | -0.0 | 0.0 | 0.0 | 0.0 |

Figure 29 First integrals. 13 mm gap. CP+ mode

Figure 30 First integrals. 20 mm gap. CP+ mode

Figure 31 First integrals. 30 mm gap. CP+ mode

Figure 32: First integrals. 40 mm gap. CP+ mode

Figure 33 First integrals. 120 mm gap. CP+ mode

Table 6: Multipoles at  increasing gaps in CP+ mode.

| Gap (mm) | Mode | Multipole | Skew | Skew error | Normal | Normal error |
| --- | --- | --- | --- | --- | --- | --- |
| 13 | CP+ | 2-pole (Gcm) | -10.3 | 1.7 | 51.4 | 2.3 |
|  |  | 4-pole (G) | 16.2 | 3.9 | -22.2 | 5.2 |
|  |  | 6-pole (G/cm) | 51.0 | 7.9 | 102.9 | 10.3 |
|  |  | 8-pole (G/cm^2) | -48.8 | 7.0 | 49.2 | 9.1 |
|  |  | 10-pole (G/cm^3) | -53.7 | 9.2 | -85.0 | 12.0 |
|  |  | 12-pole(G/cm^4) | 22.5 | 3.9 | -22.2 | 5.1 |
| Gap (mm) | Mode | Multipole | Skew | Skew error | Normal | Normal error |
| 20 | CP+ | 2-pole (Gcm) | -8.7 | 0.5 | 74.8 | 1.3 |
|  |  | 4-pole (G) | -13.6 | 0.5 | 17.8 | 1.2 |
|  |  | 6-pole (G/cm) | 5.1 | 0.5 | -9.4 | 1.3 |
|  |  | 8-pole (G/cm^2) | 2.4 | 0.1 | -3.2 | 0.3 |
|  |  | 10-pole (G/cm^3) | 0.0 | 0.1 | 0.3 | 0.2 |
| Gap (mm) | Mode | Multipole | Skew | Skew error | Normal | Normal error |
| 30 | CP+ | 2-pole (Gcm) | -7.2 | 0.2 | 68.7 | 0.2 |
|  |  | 4-pole (G) | -9.2 | 0.2 | 14.1 | 0.2 |
|  |  | 6-pole (G/cm) | 5.6 | 0.2 | -7.3 | 0.2 |
|  |  | 8-pole (G/cm^2) | 0.5 | 0.0 | -2.7 | 0.0 |
|  |  | 10-pole (G/cm^3) | -0.3 | 0.0 | 0.4 | 0.0 |
| Gap (mm) | Mode | Multipole | Skew | Skew error | Normal | Normal error |
| 40 | CP+ | 2-pole (Gcm) | -2.6 | 0.1 | 58.6 | 0.2 |
|  |  | 4-pole (G) | -6.8 | 0.1 | 9.1 | 0.2 |
|  |  | 6-pole (G/cm) | 4.8 | 0.1 | -5.4 | 0.2 |
|  |  | 8-pole (G/cm^2) | -0.4 | 0.0 | -1.9 | 0.0 |
|  |  | 10-pole (G/cm^3) | -0.3 | 0.0 | 0.3 | 0.0 |
| Gap (mm) | Mode | Multipole | Skew | Skew error | Normal | Normal error |
| 120 | HP | 2-pole (Gcm) | -1.1 | 0.1 | 18.5 | 0.1 |
| Gap (mm) | Mode | 4-pole (G) | -0.7 | 0.1 | 2.1 | 0.0 |
| 120 | CP+ | 6-pole (G/cm) | -0.1 | 0.1 | -0.2 | 0.1 |
|  |  | 8-pole (G/cm^2) | 0.0 | 0.0 | -0.1 | 0.0 |
|  |  | 10-pole (G/cm^3) | -0.0 | 0.0 | -0.0 | 0.0 |

Figure 34 First integrals. 13 mm gap. CP- mode

Figure 35 First integrals. 20 mm gap. CP- mode

Figure 36 First integrals. 30 mm gap. CP- mode

Figure 37 First integrals. 40 mm gap. CP- mode

Figure 38 First integrals. 120 mm gap. CP- mode

Table 7: Multipoles at  increasing gaps in CP- mode.

| Gap (mm) |  | Mode | Multipole | Skew | Skew error | Normal | Normal error |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 13 |  | CP- | 2-pole (Gcm) | -8.8 | 3.6 | 104.2 | -8.8 |
|  |  |  | 4-pole (G) | 12.1 | 8.2 | 14.4 | 12.1 |
|  |  |  | 6-pole (G/cm) | -94.6 | 16.4 | -38.0 | -94.6 |
|  |  |  | 8-pole (G/cm^2) | 9.3 | 14.5 | -0.0 | 9.3 |
|  |  |  | 10-pole (G/cm^3) | 89.3 | 19.1 | 5.7 | 89.3 |
|  |  |  | 12-pole(G/cm^4) | -8.8 | 8.1 | -1.0 | -8.8 |
| Gap (mm) |  | Mode | Multipole | Skew | Skew error | Normal | Normal error |
| 20 |  | CP- | 2-pole (Gcm) | -11.4 | 1.1 | 89.5 | 0.4 |
|  |  |  | 4-pole (G) | 9.4 | 1.0 | 14.8 | 0.4 |
|  |  |  | 6-pole (G/cm) | 3.6 | 1.1 | -27.8 | 0.4 |
|  |  |  | 8-pole (G/cm^2) | -0.0 | 0.2 | -3.3 | 0.1 |
|  |  |  | 10-pole (G/cm^3) | 0.8 | 0.2 | 2.7 | 0.1 |
| Gap (mm) |  | Mode | Multipole | Skew | Skew error | Normal | Normal error |
| 30 |  | CP- | 2-pole (Gcm) | -5.8 | 0.3 | 74.1 | 0.2 |
|  |  |  | 4-pole (G) | 5.4 | 0.2 | 13.9 | 0.2 |
|  |  |  | 6-pole (G/cm) | 5.5 | 0.3 | -16.4 | 0.2 |
|  |  |  | 8-pole (G/cm^2) | -0.7 | 0.1 | -2.7 | 0.0 |
|  |  |  | 10-pole (G/cm^3) | -0.1 | 0.0 | 1.4 | 0.0 |
| Gap (mm) |  | Mode | Multipole | Skew | Skew error | Normal | Normal error |
| 40 |  | CP- | 2-pole (Gcm) | 1.7 | 0.1 | 61.7 | 0.2 |
|  |  |  | 4-pole (G) | 2.1 | 0.1 | 6.5 | 0.2 |
|  |  |  | 6-pole (G/cm) | 4.7 | 0.1 | -8.6 | 0.2 |
|  |  |  | 8-pole (G/cm^2) | -1.0 | 0.0 | -1.7 | 0.0 |
|  |  |  | 10-pole (G/cm^3) | -0.3 | 0.0 | 0.6 | 0.0 |
| Gap (mm) |  | Mode | Multipole | Skew | Skew error | Normal | Normal error |
| 120 | CP- |  | 2-pole (Gcm) | 2.9 | 0.1 | 16.3 | 0.1 |
|  |  |  | 4-pole (G) | 0.2 | 0.1 | 1.9 | 0.1 |
|  |  |  | 6-pole (G/cm) | -0.3 | 0.1 | -0.3 | 0.1 |
|  |  |  | 8-pole (G/cm^2) | -0.1 | 0.0 | -0.2 | 0.0 |
|  |  |  | 10-pole (G/cm^3) | 0.0 | 0.0 | -0.0 | 0.0 |

Figure 39 First integrals. 13 mm gap. VP mode

Figure 40 First integrals. 20 mm gap. VP mode

Figure 41 First integrals. 30 mm gap. VP mode

Figure 42 First integrals. 40 mm gap. VP mode

Figure 43 First integrals. 40 mm gap. VP mode

Table 8 Multipoles at  increasing gaps in CP- mode.

| Gap (mm) |  | Mode | Multipole | Skew | Skew error | Normal | Normal error |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 13 |  | VP | 2-pole (Gcm) | -14.0 | 1.9 | 60.5 | 1.5 |
|  |  |  | 4-pole (G) | -23.6 | 4.4 | 99.6 | 3.3 |
|  |  |  | 6-pole (G/cm) | -12.6 | 8.8 | 44.1 | 6.6 |
|  |  |  | 8-pole (G/cm^2) | -8.6 | 7.8 | -37.3 | 5.9 |
|  |  |  | 10-pole (G/cm^3) | 7.7 | 10.3 | -42.6 | 7.7 |
|  |  |  | 12-pole(G/cm^4) | 6.6 | 4.4 | 10.9 | 3.3 |
| Gap (mm) |  | Mode | Multipole | Skew | Skew error | Normal | Normal error |
| 20 |  | VP | 2-pole (Gcm) | -21.6 | 0.6 | 59.6 | 1.1 |
|  |  |  | 4-pole (G) | -15.5 | 0.5 | 47.0 | 1.0 |
|  |  |  | 6-pole (G/cm) | 10.8 | 0.5 | -6.0 | 1.1 |
|  |  |  | 8-pole (G/cm^2) | 2.8 | 0.1 | -6.8 | 0.2 |
|  |  |  | 10-pole (G/cm^3) | -0.4 | 0.1 | 0.2 | 0.2 |
| Gap (mm) |  | Mode | Multipole | Skew | Skew error | Normal | Normal error |
| 30 |  | VP | 2-pole (Gcm) | -9.2 | 0.1 | 53.1 | 0.2 |
|  |  |  | 4-pole (G) | -9.0 | 0.2 | 24.1 | 0.4 |
|  |  |  | 6-pole (G/cm) | 8.1 | 0.1 | -1.7 | 0.2 |
|  |  |  | 8-pole (G/cm^2) | 1.7 | 0.1 | -5.7 | 0.2 |
|  |  |  | 10-pole (G/cm^3) | -0.5 | 0.0 | -0.0 | 0.0 |
| Gap (mm) |  | Mode | Multipole | Skew | Skew error | Normal | Normal error |
| 40 |  | VP | 2-pole (Gcm) | -2.0 | 0.1 | 51.9 | 0.2 |
|  |  |  | 4-pole (G) | -5.9 | 0.1 | 10.3 | 0.2 |
|  |  |  | 6-pole (G/cm) | 4.9 | 0.1 | -1.4 | 0.2 |
|  |  |  | 8-pole (G/cm^2) | -0.3 | 0.0 | -2.0 | 0.0 |
|  |  |  | 10-pole (G/cm^3) | -0.3 | 0.0 | -0.0 | 0.0 |
| Gap (mm) |  | Mode | Multipole | Skew | Skew error | Normal | Normal error |
| 120 | VP |  | 2-pole (Gcm) | -2.2 | 0.1 | 18.9 | 0.0 |
|  |  |  | 4-pole (G) | -0.8 | 0.1 | 1.1 | 0.0 |
|  |  |  | 6-pole (G/cm) | 0.1 | 0.1 | -0.0 | 0.0 |
|  |  |  | 8-pole (G/cm^2) | 0.1 | 0.0 | 0.0 | 0.0 |
|  |  |  | 10-pole (G/cm^3) | -0.0 | 0.0 | -0.0 | 0.0 |

## First derivative of the first integrals

The first derivative of the first integrals has been computed from the polynomial fit of the first integral, where a sufficient high number of polynomial terms has been used, to achieve convergence.

Figure 44 First derivative of the first integral.13 mm gap, HP Mode.

Figure 45 First derivative of the first integral.20 mm gap, HP Mode.

Figure 46 First derivative of the first integral.30 mm gap, HP Mode.

Figure 47 First derivative of the first integral.40 mm gap, HP Mode.

Figure 48 First derivative of the first integral.120 mm gap, HP Mode.

Figure 49 First derivative of the first integral.13 mm gap, CP+ Mode.

Figure 50 First derivative of the first integral.20 mm gap, CP+ Mode.

Figure 51 First derivative of the first integral.30 mm gap, CP+ Mode.

Figure 52 First derivative of the first integral.40 mm gap, CP+ Mode.

Figure 53 First derivative of the first integral.40 mm gap, CP+ Mode.

Figure 54 First derivative of the first integral.13 mm gap, CP- Mode.

Figure 55 First derivative of the first integral.20 mm gap, CP- Mode.

Figure 56 First derivative of the first integral.30 mm gap, CP- Mode.

Figure 57 First derivative of the first integral.40 mm gap, CP- Mode.

Figure 58 First derivative of the first integral.120 mm gap, CP- Mode.

Figure 59 First derivative of the first integral.13 mm gap, VP Mode.

Figure 60 First derivative of the first integral.20 mm gap, VP Mode.

Figure 61 First derivative of the first integral.30 mm gap, VP Mode.

Figure 62 First derivative of the first integral.40 mm gap, VP Mode.

Figure 63 First derivative of the first integral.120 mm gap, VP Mode.

## The second integrals

The second integrals were measured by performing both twisted and normal flip coil turn measurements at all gaps. The second integrals Jx and Jy can be calculated from the measurements with twisted flip coil and the measurements of the first integrals

A normal and a twisted flip quick coil scan have been measured in the y=0 plane over the interval –2.5  x  2.5 cm with 0.2 cm steps at each gap. The data analysis is done with an Igor Pro macro that loads the first integral and second integral data files, converts the length unit to centimeters and calculates the average second integrals at each x-position. Our standard procedure is to correct the first and second field integrals by subtracting of the ambient field integral contribution. This was chosen to be at open gap (120 mm).  In the CP- mode, at 13 mm gap, the horizontal second field integral, is just outside the specification.

Figure 64 Second integrals. 13 mm gap. HP mode

Figure 65 Second integrals. 20 mm gap. HP mode

Figure 66  Second integrals. 30 mm gap. HP mode

Figure 67 Second integrals. 30 mm gap. HP mode

Figure 68 Second integrals. 120 mm gap. HP mode

Figure 69 Second integrals. 13 mm gap. CP+ mode

Figure 70 Second integrals. 20 mm gap. CP+ mode

Figure 71 Second integrals. 30 mm gap. CP+ mode

Figure 72 Second integrals. 40 mm gap. CP+ mode

Figure 73 Second integrals. 120 mm gap. CP+ mode

Figure 74 Second integrals. 13 mm gap. CP- mode

Figure 75 Second integrals. 20 mm gap. CP- mode

Figure 76 Second integrals. 30 mm gap. CP- mode

Figure 77 Second integrals. 40 mm gap. CP- mode

Figure 78 Second integrals. 120 mm gap. CP- mode

Figure 79 Second integrals. 13 mm gap. VP mode

Figure 80 Second integrals 20 mm gap. VP mode

Figure 81 Second integrals 30 mm gap. VP mode

Figure 82 Second integrals 40 mm gap. VP mode

Figure 83 Second integrals 40 mm gap. VP mode

# Magnetic Measurements with tunecorrectors

In this section, magnetic measurements are presented for the device, after ESRF-style tune correctors have been installed.  4 sets of shims were installed, 16 individual shims in all.

In this section we will present integral and hall scans for the 4 interesting phases, at minimum gap, as well as flip-coil measurements at the same phases.

The resultant first integrals, defined as the sum of the measured, and dynamic integrals, was optimized to be smallest in the circular phases at minimum gap.  In the HP mode, the effect of the shims actually worsens the resultant integral.  In the VP mode, the shims also reduce the resultant dynamic integrals.  At 30 mm gap, the HP mode and VP mode are manageable.

As expected, there was no drastic effect on the undulator phase error, nor trajectory straightness.

## First integral scans after tune shim installation

### 13 mm gap

Figure 84 First integrals at 13 mm gap. HP mode.  Also shown is the calculated dynamic effect, and the sum of the measured and dynamic effect.

Figure 85 First integrals at 13 mm gap. CP+ mode.  Also shown is the calculated dynamic effect, and the sum of the measured and dynamic effect.

Figure 86 First integrals at 13 mm gap. CP+ mode.  Also shown is the calculated dynamic effect, and the sum of the measured and dynamic effect.

Figure 87 First integrals at 13 mm gap. VP mode.  Also shown is the calculated dynamic effect, and the sum of the measured and dynamic effect.

### 30 mm gap

Figure 88 First integrals at 30 mm gap. HP mode.  Also shown is the calculated dynamic effect, and the sum of the measured and dynamic effect.

Figure 89 First integrals at 30 mm gap. CP+ mode.  Also shown is the calculated dynamic effect, and the sum of the measured and dynamic effect.

Figure 90 First integrals at 30 mm gap. CP- mode.  Also shown is the calculated dynamic effect, and the sum of the measured and dynamic effect.

Figure 91 First integrals at 30 mm gap. VP mode.  Also shown is the calculated dynamic effect, and the sum of the measured and dynamic effect.

## Hall probe scans on the undulator axis(13 mm gap)

### HP Mode

Figure 92 The electron orbits for a 3 GeV electron at 13 mm gap for x = y = 0(HP Mode).. The horizontal orbit averaged over one undulator period is also shown. After tuneshim installation

### CP+ Mode

Figure 93 The electron orbits for a 3 GeV electron at 13 mm gap for x = y = 0(CP+ Mode).. The horizontal orbit averaged over one undulator period is also shown.

CP- Mode

Figure 94 The electron orbits for a 3 GeV electron at 13 mm gap for x = y = 0(CP- Mode).. The horizontal orbit averaged over one undulator period is also shown.

## Analysis of peak field rolloff

For this device, there was some emphasis on the peak field rolloff performance.  During the conceptual design stage, we optimized the exact block dimensions, to improve the derivative in the field roll off, from 500G/cm to below 400 G/cm.  The field roll off was evaluated by locally scanning over several poles, in the x-direction.

### CP mode(13 mm gap)

Figure 95 Field roll off

Figure 96 First derivative of field roll-off.

Figure 97 Second derivative of field roll-off.

## Reproducibility of the magnetic array

In this section, we present some investigations of the reproducilibity of the magnetic array, upon the phasing motion.  As is clear from the conceptual design report, the magnetic force vectors for this structure are very large, the force vectors change direction upon phasing.

We have investigated, at minimum gap, the reproducibility of the circular polarization mode, after travelling to the HP mode, and back again. We can study the average magnetic taper and see if it changes at all.  This can be done for both the vertical and horizontal field, based on the value of the peak field at each pole.  We see little change in both the average peak field as well as the average magnetic taper.

Figure 98 Magnetic taper fit, based on vertical peak fields measured at each pole

Figure 99 Magnetic taper fit, based on horizontal peak fields measured at each pole

Table 9 Summary of horizontal peak field taper upon phasing from CP+

| ### Scan number | ### Motion | ### Av.Peak field | ### Av. Taper(/pole) |
| --- | --- | --- | --- |
| ### 1 | ### CP+ to HP to CP+ | ### 0.73245 | ### -9.925e-005 |
| ### 2 | ### CP+ to HP to CP+ | ### 0.73248 | ### -9.9445e-005 |
| ### 3 | ### CP+ to HP to CP+ | ### 0.73217 | ### -9.8275e-005 |
| ### 4 | ### CP+ to HP to CP- to CP+ | ### 0.73206 | ### -9.4014e-005 |

Table 10 Summary of vertical peak field taper upon phasing from CP+

| ### Scan number | ### Motion | ### Av.Peak field | ### Av. Taper |
| --- | --- | --- | --- |
| ### 1 | ### CP+ to HP to CP+ | ### 0.73304 | ### -0.00011778 |
| ### 2 | ### CP+ to HP to CP+ | ### 0.7331 | ### -0.00011969 |
| ### 3 | ### CP+ to HP to CP+ | ### 0.73285 | ### -0.00011733 |
| ### 4 | ### CP+ to HP to CP- to CP+ | ### 0.73305 | ### -0.00011491 |

Table 11 Summary of horizontal peak field taper upon phasing from CP-

| ### Scan number | ### Motion | ### Av.Peak field | ### Av. Taper(/pole) |
| --- | --- | --- | --- |
| ### 1 | ### CP- to HP to CP- | ### 0.73459 | ### -0.00030956 |
| ### 2 | ### CP- to HP to CP- | ### 0.73463 | ### -0.0003103 |
| ### 3 | ### CP- to HP to CP- | ### 0.73448 | ### -0.00030806 |

Table 12 Summary of vertical peak field taper upon phasing from CP-

| ### Scan number | ### Motion | ### Av.Peak field | ### Av. Taper(/pole) |
| --- | --- | --- | --- |
| ### 1 | ### CP- to HP to CP- | ### 0.72996 | ### 0.00019122 |
| ### 2 | ### CP- to HP to CP- | ### 0.73013 | ### 0.0001864 |
| ### 3 | ### CP- to HP to CP- | ### 0.73002 | ### 0.00018925 |

We also examined the effect on the phase error at 30 mm gap, when changing gaps from 13 mm to 30 mm, as compared to changing gaps from 40 mm to 30 mm. The phase error, in each case, is shown in Figure 100, and summarized in Table 13.  We see no discernible change in the magnetic results which could indicate mechanical hysteresis of any kind.

Figure 100 Phase error, in the VP mode, at 30 mm gap, upon gap change from 13 mm and 40 mm.

Table 13 Summary of vertical peak field taper upon changing gap( VP mode)

| ### Scan number | ### Motion | ### Phase error | ### Av. Peak field |
| --- | --- | --- | --- |
| ### 1 | ### 13 mm to 30 mm | ### 3.84 | ### 0.474168 |
| ### 2 | ### 40 mm to 30 mm | ### 3.83 | ### 0.474175 |

## Correction coils HP mode

We have calibrated the coils in the HP mode at different gaps.  Due to the close to unit permeability of the magnet material, we expect little effect of moving the undulator phase.

There are 12 vertical windings and 8 horizontal windings.   We see little multipole content coming from the coils.  The signal coming from the horizontal coil is quite weak, which means that the uncertainty, at the smallest gap, becomes comparable to the signal itself.

### Vertical coils

Figure 101 Second vertical integral data when upstream and downstream coils are powered with +10 A

Figure 102 Second vertical integral data when upstream and downstream coils are powered with -10 A

### Horizontal coils

Figure 103 Second horizontal integral data when upstream and downstream coils are powered with +10 A

Figure 104 Second horizontal integral data when upstream and downstream coils are powered with +10 A

# General Comments

## Mounting of the permanent magnets and the shimming process

The magnetic blocks mounted on the device, were received from Arnold Magnetics.  They were all within the magnetic specifications, as presented by Arnold Magnetics using a Helmholtz system. In order to gain the strongest possible magnets, we chose isostatically pressed magnets, as opposed to transverse die pressed magnets.  In this manner, Arnold Magnetics were able to provide magnets with a minimum closed circuit remanence above the specified 1.06 T.

Figure 105 Remanence data obtained from Arnold magnetics

Figure 106 Remanence data from Arnold magnetics

Due to our 3+1 keeper design, the magnets had to be divided into several different populations, prior to being mounted into the keepers.  The magnets for the B-keepers, which consist of two opposing vertical magnets, along with a horizontal block, were chosen such that the magnetic misorientation of the vertical magnets were minimized.  Care was also taken to ensure that the remanence of the magnets in the module were similar, in order to minimize the observed phase error.

The field integral contributions of the A and B modules were all measured for -50  x  50 mm in steps of 1 mm with the center of the flip coil 6.5 mm from the magnets.  These field integral contributions are loaded into a database.  This database takes care of block rotations and block flips such that the measured data is modified according to position on the undulator. Sample integral contributions are shown in Figure 107.

Figure 107 Sample field integral contributions shown for different magnet blocks.

The first period was mounted on the girder starting on the upper girder at one end. After mounting one period the field integrals were measured 6.5 mm from the magnets and the next period, consisting of an A+B- module set, was selected in order to minimize the field integral variations. The Hall bench was used to ensure accurate positioning and alignment of each module.  Finally both end sections were selected and mounted.

The device mas shimmed mainly by shimming magnetic poles in B-modules,as well as a little module swapping.  The swapping and shimming process consisted initially of trajectory straightening, and then of multipole corrections.  Magic finger corrections were applied to minimize the integral multipoles. These, however, do not correct phase dependent integrals.

## Conclusion

We have here presented the final results of the magnetic measurement test of the SSRL 140 mm Apple-II undulator.  The details of the measurement setup have been described and the reproducibility of the Hall probe bench and the flip coil bench have been documented. From the Hall probe x and y-scans it has been documented that the undulator was well aligned.

The magnetic on-axis properties have been measured as a function of the gap. It was found that the undulator period is 139.98  mm.  The electron straightness and electron phase error is quite good, as function of gap and phase.  Unfortunately, the undulator does not meet the field specification which concerns the minimum photon energy. We designed the device to have a minimum photon energy of 6.7 eV, in the circular polarization phase, but ended up with 7.4 eV, corresponding to several percent of peak field loss.  We have performed detailed simulations, to check if, in fact, there was some risk of demagnetization, in one of the phases but this was not found to be the case.  The calibration of the probe, has also been checked against a commercial Group3 teslameter, which indicated that the calibration of the vertical field was fine.  We also experience a phase dependent multipole effect, which means that we are slightly off-spec in the integral specification, due to a rapid field fluctuation near the undulator axis.

It was found that tune correctors could be implemented with minimal effect on the skew integrals, and small effect on the electron trajectory, and electron phase error.