## Magnetic Measurement Traveler -  Dipole

| Item | Result | Form/Unit |
| --- | --- | --- |
| Project | E163 | text |
| Magnet Type: | Dipole | Dipole/Quad |
| Magnet Name: | 1.13D6B | text |
| Magnet Name Label on Magnet: | B | text |
| Magnet Main polarity: | Top Pole is North |  |
| Vendor serial number from magnet label: | N/A | # |
| SLAC barcode number: | N/A | # |
| SLAC drawing number: | SK-235-978-31 | # |
| Beam direction arrows – both sides? | N | Y/N |
| Polarity designators affixed? | Y | Y/N |
| SLAC approved electrical safety covers? | N | Y/N |
| SLAC approved lifting eyes? | Y | Y/N |
| Cooling water pressure drop: | 120 | psi |
| Cooling water flow: | 0.56 | gpm |
| Cooling water temperature input: | 26.8 | oC |
| Cooling water temp rise at max current: | 46.70 | oC |
| Asymptotoic coil T at max current : | 46.68 @ 1.5 hrs | oC(@ ? hrs) |
| Resistance Main | 0.044 | Ohm |
| Resistance Trim | 0.093 | Ohm |
| Inductance Main | 2.73 | mH @ 100 Hz |
| Inductance Trim | 0.647 | mH @ 100 Hz |
| Date of arrival to MM group: | 11-01-2005 | mm/dd/yy |
| Responsible measurement operator name: | Scott Anderson | print |
| Responsible measurement operator signature: |  | script |
| Date of completion of SLAC tests: | 05/02/2006 | mm/dd/yy |
| Magnet Photo | Y | Y/N |
| Magnet CMM Data | Y | Y/N |

State "N/A" for items not applicable.  State "see note # below" for comments appended.

See appendix below for magnet polarity control protocol.

Location of on-line magnetic measurement data:

http://www-group.slac.stanford.edu/met/MagMeas/MAGDATA/E163/Dipole/1.13D6B/

Rev 1

Traveler Version Number:

| Stretched Wire (Y/N) | Y | Rotating Coil (Y/N) | N |
| --- | --- | --- | --- |
| Hall Probe Map (Y/N) | Y | Temperature (Y/N) | Y |

Measurement type:

| CoilName: |  | Radius: |  | mm |
| --- | --- | --- | --- | --- |

Rotating coil:

| N/A | Minutes @ |  | Amps |
| --- | --- | --- | --- |

Warmup before measurements:

Standardization Cycles Performed before measurements:

| Number of cycles | 3 | Minimum Current (A) | 240 |
| --- | --- | --- | --- |
| Ramp Type | 3 Linear | Maximum Current (A) | 0 |
| Ramp Rate (A/s) | 20 | Settle Time @ Max and Minimum Current (sec) | 2 |

Dipole Magnets- Stretched Wire the incoming x =0 line (1 inch off center).

| 240.564 | Amps @ | 0.218417 | T-m |
| --- | --- | --- | --- |

Current at maximum integrated field:

| 120.273 | Amps @ | 0.131270 | T-m |
| --- | --- | --- | --- |

Current at nominal integrated field:

Coefficients of Magnetic Moment Series (expressed in % of principal moment):

| Dipole n = 1 | Quadrupole n = 2 | Sextupole n = 3 | Decapole n = 5 |
| --- | --- | --- | --- |
| 100% |  |  |  |

| Dipole Magnet Hall Probe |
| --- |
| Hall Probe Type | Group3 | Hall Probe # | 01231690 |

| 240.2433 | Amps @ | -0.855550 | T |
| --- | --- | --- | --- |

Current at maximum poletip field:

| 120.0647 | Amps @ | -0.494217 | T |
| --- | --- | --- | --- |

Current at nominal poletip field:

Report and Notes: