Magnet Measurement Plan

| Logistic Info |
| --- |
| Person Responsible | Jim Spencer |
| Phone # | x3385 |
| Start Date | 7/12/2006 |
| Finish Date |  |
| Acct # for alignment |  |

| Magnet Designation and Type |
| --- |
| Project Name | E163 |
| Magnet Type | Dipole |
| Magnet Name/Number | Spectrometer |
| Serial # / PC # |  |
| Trim Coils | No | Trim Types |  |
| Magnet gap or diameter (mm) |  |

| Alignment |
| --- |
| Coarse (2.0 mm)/Precision (0.05 mm) |
| Has the magnet been fiducalized | Yes |
| If above is No , what is origin point of the magnet, what is its coordinate system and what surfaces should be used for alignment? (Please provide drawing) | 0,0 point (the vertex point) is the point defined by the intersection of the lines perpendicular and at the center of the entrance and exit poles |

| Special Alignment Tolerances |
| --- |
| X |  |
| Y |  |
| Pitch |  |
| Yaw |  |
| Roll |  |

| Water and Power |
| --- |
| Water Pressure | 120 | PSI |
| Water Flow | 3.6 | Gpm |
| Maximum Current Main | 250 | Amps |
| Max Current Trims | N/A | Amps |

| Polarity |
| --- |
| Main Coil Polarity. | North pole on the top. |
| Dipole Trim Polarity | N/A |
| Quad Trim Polarity | N/A |

| Warmup and Standardization |
| --- |
| Magnet warm up. Time and Current | ? mins @ ? Amps |
| Number of Stdz cycles | 3 |
| Ramp Rate (A/s) | 10 | Minimum Current (A) | 0 |
| Settle Time (sec) | 60 | Maximum Current (A) | 250 |
| Ramp Type (Linear or Three Linear or Cosine | Three Linear |

| Temperature Tests |
| --- |
| Temperature and Resistance Test | Yes |
| Temp Items to be measured (Return water, coils ?) | Return Water, Top Coil |
| Currents (Amps) | 100, 150, 200, 250 |
| Check water flow | Yes – 3.6 gpm |

| Integrated Strength |
| --- |
| Main | Yes | Currents | Using hall probe at x = 0, z = -0.3 m .0-250 A in 20 Amps steps (Note 1) |
| Dipole Trim | No | Currents |  |
| Quad Trim | No | Currents |  |

| Harmonics |
| --- |
| Main | No | Currents |  |
| Dipole Trim | No | Currents |  |
| Quad Trim | No | Currents |  |

| Harmonics Tolerances (as % of main harmonic, i.e. main = 100%) |
| --- |
|  | Dipole n = 1 | Quadrupole n = 2 | Sextupole n = 3 | Other n = |
| Main |  |  |  |  |
| Dipole Trim |  |  |  |  |
| Quad Trim |  |  |  |  |

| Field Map |
| --- |
| Type of map? (Transverse, axial, 3D) | By vs z and x at the exit pole. (0,0) is the vertex point. |
| Dimensions ( X, Y, Z) | X: -0.05 to +0.05 Z -0.58 to +0.10 m |
| Spacing between step (X, Y, Z) | X: 0.005 m Z 0.005 m |

Notes and/or Additional Tests:

- Use Hall probe in pole middle to what currents give 0.38, .40 and .42 T.

  - By = .38 T    Imag = 157.2076 Amps

  - 0.4T    165.6048 Amps

  - 0.42 T  174.0169 Amps

  - See By vs Imag fit for run 2.bmp for fit.  Fit was 5th order polynomial