## SLAC Traveler for SXR Undulator #29A804-000 in MMF

Dataset 002

This traveler is intended to cover preparation, magnetic measurements, and fiducialization of the SXR undulator segments at MMF.

Preparation:

- Remove all protective covers. Install wooden platform between the undulator and granite.

- Remove plastic cover. Clean up the granite.

- Attach temperature sensors.

- Check linear potentiometers. Install brackets if necessary.

- Install coil supports.

| Vendor serial number from undulator label: |  |
| --- | --- |
| Technician (initials): |  |
| Date/Time: |  |

- Call controls group to hook-up and check the control system. Check completed.

- Call alignment crew to align the undulator to the bench. Alignment done.

- Open gap to 10mm (tuning gap). Check gap straightness with the touch probe.

- Correct pitch and yaw, if necessary.

- Measure mechanical axis shift vs. Gap. Compare to the previous measurements.

- Check undulator pitch vs Gap with the touch probe.

- Remove touch probe. Install the Hall probe, align.

| Engineer(initials)/date and time: |  |
| --- | --- |

Measurement process checklist:

| Engineer (initials): |  |
| --- | --- |
| Measurements started, Date (mm-dd-yyyy): |  |
| Align the undulator magnetically at tuning Gap = 10mm; |
| Trajectories, phase errors and phase matching checked and within the tolerances at commissioning gap. |
| Kugler bench OFF |
| Long coil set, PDI and XPS ON |
| Field integrals checked. |
| Final scan done, coil |
| Kugler bench ON and re-initialized. |
| Reference magnet measured |
| Voltmeters checked |
| MA position vs Y checked. |
| Final scan done, Hall probe |
| PMs measured for fiducialization, file with results created. |
| Fiducialization done |
| Capacitive sensor measurements done. |
| Vacuum chamber limit switches checked. |
| Controller Configuration saved. |

The following information is to be noted by an engineer:

| Average X (m): |  |
| --- | --- |
| Average Y (m): |  |
| Final Pitch (µrad): |  |
| Final Yaw (µrad): |  |
| Magnetic axis position X (m): |  |
| Magnetic axis position Y (m): |  |
| Coil X-position (m) |  |
| Coil Y-position (m) |  |
| Commissioning Gap (mm) |  |
| Reference magnet measurement (T) |  |
| Voltmeter X (1V) 1.00002 |  |
| Voltmeter Y (1V) 1.00002 |  |
| Voltmeter Y (6V) 6.00003 |  |
| Voltmeter Y (-6V) -6.00004 |  |

Fiducialization:

Upon completion of the tuning the undulator should be fiducialized.

Measure position of the US and DS PM centers w.r.t. the undulator center line.

The following information is to be noted by an engineer:

| Engineer (Initials): |  |
| --- | --- |
| Upstream RM X offset (m): |  |
| Upstream RM Y offset (m): |  |
| Downstream RM X offset (m): |  |
| Downstream RM Y offset (m): |  |

Have an alignment crew to measure offsets between RM and the undulator fiducials.

| Surveyor (initials): |  |
| --- | --- |
| Date (mm/dd/yyyy): |  |

Prepare the undulator for the move into the storage area:

- Power OFF the controller. Disconnect the network cable.

- Remove linear potentiometer brackets, if installed.

- Remove long coil brackets.

- Remove temperature sensors.

- Install all protective covers. Remove wooden platform between the undulator and granite.

- The undulator is ready for move into the storage area.

| Technician (initials) / date: |  |
| --- | --- |

URL for data:

V:\MET\MagServe\MagData\LCLS-II\Undulator\HE_SXU_000 \DATASET0002 \Final Results\

| Undulator check completed (signed): |  |
| --- | --- |
| Date (mm/dd/yy) |  |