## SLAC Traveler for SXR Undulator #29A804-003 in MMF

This traveler is intended to cover preparation, magnetic measurements, and fiducialization of the SXR undulator segments at MMF.

Preparation:

- Remove all protective covers. Install wooden platform between the undulator and granite.

The following information has to be noted upon moving an undulator into the lab by SLAC technician:

| Technician (initials): |  |
| --- | --- |
| Date (mm-dd-yyyy): |  |
| Vendor serial number from undulator label: |  |

- Call alignment crew to align the undulator to the bench.

- Attach temperature sensors; 3 – to the top & 2 – to the bottom jaws.

- Call controls group to hook-up and check the control system.

- Make sure all emergency stop buttons are de-activated.

- Install linear potentiometer brackets.

- Check torque of all magnet motion screws with 4.4 inch-lbs torque wrench. Report loose screws, if any.

- Label modules from 1 to 5.

- Check spacing between poles and magnets with 0.015” brass shim. Correct if necessary.

| Technician (initials): |  |
| --- | --- |

- Open gap to 10mm (tuning gap). Check gap straightness with the touch probe.

- Correct encoders’ offsets to make jaws parallel and level.

- Measure mechanical axis shift vs. Gap. Fit parabola to the measurements. Input new coefficients into “GapOffsetCoefficients.txt” file.

- Set pitch limit switches.

| Engineer(initials): |  |
| --- | --- |

Tuning.

- Remove touch probe. Install the Hall probe, align.

Follow the check list below, consult with the tuning test plan (LCLS-TN-17-nn ).

Tuning process checklist:

| Engineer (initials): |  |
| --- | --- |
| Date (mm-dd-yyyy): |  |
| Align the undulator magnetically at tuning Gap = 10mm; adjust encoder offsets if necessary, , record the following: |
| Fields measured for tuning trajectories |
| Upstream PM measured |  |
| Kugler bench OFF |  |
| Long coil set, PDI and XPS ON |  |
| Optical measurement of coil position done |  |
| Field integrals measured for tuning trajectories |  |
| Coil and stages in park position, PDI and XPS OFF |  |
| Shim table prepared |  |
| Gap open to 180mm |  |
| Undulator trajectory shimming completed |  |
| Kugler bench ON and re-initialized |  |
| Trajectories and phase errors checked. If additional shimming needed, repeat Hall probe measurements and tuning. |  |
| Trajectories and phase errors tuned to tolerances |  |
| Fields measured for tuning ends |  |
| Kugler bench OFF |  |
| Long coil set, PDI and XPS ON |  |
| Field integrals measured for tuning first and second field integrals |  |
| Coil and stages in park position, PDI and XPS OFF |  |
| Shim table prepared |  |
| Gap open to 180mm |  |
| Undulator field integral shimming completed. |  |
| Gap closed to 10mm. Field integrals checked. If additional shimming needed, repeat coil measurements and tuning. |  |
| Kugler bench ON and re-initialized |  |
| Trajectories and phase errors checked and within the tolerances. |  |
| Reference magnet measured |  |
| Voltmeters checked |  |
| Final scan done, coil |  |
| Final scan done, Hall probe |  |
| Gap clearance checked |  |
| Capacitive sensor measurements done |  |

The following information is to be noted by an engineer:

| Average X (m): |  |
| --- | --- |
| Average Y (m): |  |
| Final Pitch (rad): |  |
| Final Yaw (rad): |  |
| Magnetic axis position X (m): |  |
| Magnetic axis position Y (m): |  |
| Coil X-position (m) |  |
| Coil Y-position (m) |  |
| Reference magnet measurement (T) |  |
| Voltmeter X (1V) 0.999987 |  |
| Voltmeter Y (1V) 0.999987 |  |
| Voltmeter Y (6V) 5.999900 |  |
| Voltmeter Y (-6V) -5.999910 |  |

Fiducialization:

Upon completion of the tuning the undulator should be fiducialized.

Measure position of the US and DS PM centers w.r.t. the undulator center line.

The following information is to be noted by an engineer:

| Engineer (Initials): |  |
| --- | --- |
| Upstream RM X offset (m): |  |
| Upstream RM Y offset (m): |  |
| Downstream RM X offset (m): |  |
| Downstream RM Y offset (m): |  |

Have an alignment crew to measure offsets between RM and the undulator fiducials.

| Surveyor (initials): |  |
| --- | --- |
| Date (mm/dd/yyyy): |  |

Prepare the undulator for the move into the storage area:

- Power OFF the controller. Disconnect the network cable.

- Remove linear potentiometer brackets.

- Remove long coil brackets.

- Install all protective covers. Remove wooden platform between the undulator and granite.

- The undulator is ready for move into the storage area.

| Technician (initials): |  |
| --- | --- |

URL for data:

V:\MET\MagServe\MagData\LCLS-II\Undulator\SXR_0 \DATASET00 \Final Results\

| Undulator tuning completed (signed): |  |
| --- | --- |
| Date (mm/dd/yy) |  |