## Magnetic Measurement Traveler - Solenoid

| Item | Result | Form/Unit |
| --- | --- | --- |
| Project | LCLS | text |
| Magnet Type: | Solenoid | text |
| Magnet Name: | Stangenes 2 (prototype #2) | Dipole/Quad etc |
| Magnet Name Label on Magnet: | N/A | text |
| Magnet polarity: |  | text |
| Vendor serial number from magnet label: |  | # |
| SLAC barcode number: | N/A | # |
| SLAC drawing number: | N/A | # |
| Beam direction arrows – both sides? | N/A | Y/N |
| Polarity designators affixed? | Y | Y/N |
| SLAC approved electrical safety covers? | N | Y/N |
| SLAC approved lifting eyes? | Y | Y/N |
| Cooling water pressure drop: | 110 | psi |
| Cooling water flow: | 1.6 | gpm |
| Cooling water temperature input: | 22 | oC |
| Cooling water temp rise at max current: |  | oC |
| Asymptotoic coil T at max current : | 71 | oC(@ 2.5 hrs) |
| Date of arrival to MM group: | 2/28/2006 | mm/dd/yy |
| Responsible measurement operator name: | Scott Anderson | print |
| Responsible measurement operator signature: |  | script |
| Date of completion of SLAC tests: |  | mm/dd/yy |
| Magnet Photo | Y | Y/N |
| Magnet CMM Data | N | Y/N |
| Aligned | Y | Y/N |

State "N/A" for items not applicable.  State "see note # below" for comments appended.

See appendix below for magnet polarity control protocol.

Location of on-line magnetic measurement data:

http://www-group.slac.stanford.edu/met/MagMeas/MAGDATA/LCLS/Solenoid/Stangenes_2/

Rev 1

Traveler Version Number:

| Short Coil (Y/N) | Y | Long Coil (Y/N) | N |
| --- | --- | --- | --- |
| Hall Axial Map (Y/N) | ? | Temperature (Y/N) | Y |
| Hall 3-D Map (Y/N) |  |  |  |

Measurement type:

| Short Coil Name: | Coil 2 | Radius: | 0.0467 | m |
| --- | --- | --- | --- | --- |
| Long Coil Name |  | Radius: |  | m |

Rotating coil:

Warmup before measurements:

| 10 | Minutes @ | 30 | Amps |
| --- | --- | --- | --- |

Standardization Cycles Performed before measurements:

| Number of cycles | 4 | Minimum Current (A) | 0 |
| --- | --- | --- | --- |
| Ramp Type | 3 Linear | Maximum Current (A) | 30 |
| Ramp Rate (A/s) | 4 | Settle Time @ Max and Min Current (sec) | 10 |

Solenoid Magnets- Long Coil Measurements

|  | Amps @ |  | T-m |
| --- | --- | --- | --- |

Current at max integrated gradient:

|  | Amps @ |  | T-m |
| --- | --- | --- | --- |

Current at nominal integrated gradient:

Coefficients of Magnetic Moment Series (expressed in % of principal moment):

| Dipole, n = 1 | Quadrupole, n = 2 | Sextupole, n = 3 | Decapole, n = 5 |
| --- | --- | --- | --- |
|  | 100% |  |  |

Position of magnetic axis (dipole = 0) relative to CMM measured mechanical axis:

See CMM data sheet for coordinate and sign definitions.

| x = |  | mm | y = |  | mm |
| --- | --- | --- | --- | --- | --- |

Solenoid Magnets- Short Coil Measurements

Position of magnetic axis (dipole = 0) relative to alignment axis:

| x = | 0.075 | mm | y = | 0.125 | mm |
| --- | --- | --- | --- | --- | --- |

Magnetic Harmonic Moments (in G):

| Location Name z (m) | Dipole, n=1 (G) | Quad, n=2, (G) | Sext, n = 3, (G) | Oct, n = 4, (G) |
| --- | --- | --- | --- | --- |
| Entrance z = -0.0007 | 1.1 | 0.071 | 0.6 | 0.059 |
| Middle z = 0.1795 | 0.34 | 0.15 | 0.67 | 0.1 |
| Exit z = 0.3297 | 2.1 | 0.034 | 0.2 | 0.05 |

Hall Axial.

| Max Axial Field (T) | ~0.218 | Location in Z (m) | ~0.17 | Current | 30.010 Amps |
| --- | --- | --- | --- | --- | --- |

Report and Notes:

- The magnet was aligned to the inside G10 surface.   Z = 0 is the outer endg of the circular plate on the non-power lead side.  .   For the magnetic alignment, the dipole moment couldn’t be eliminated at the solenoid ends.  This means that the two circular plates at the ends are not aligned with respect to the G10. The probe was set to compromise the dipole reading at the two ends.  This gives a 3 mil offset (0.075 mm) in x and a 5 mil offset (0.125 mm) in y from the G10 alignment

- Hall Axial measurement was a hand held measurement.  All are approximate.