## SLAC Magnetic Measurement Plan and Traveler for the BROOKHAVEN Dipole Magnet LBTDPL4800#3

## Account Number to be charged: 730100

This traveler covers the magnetic measurements plan for one water-cooled dipole called LBTDPL4800#3, to be used in the linac to booster transport line of the NSLSII at Brookhaven National Lab. We do not expect to measure any other of this style dipole this year.

This is a brand new dipole made in 2011 at an outside-SLAC magnet vendor. Consequently it has no residual radioactivity. It has a laminated steel core, water cooled main coils and solid wire trim coils [which were potted with the main coils and so are not visible].

This magnet itself weighs about 1350 lb, and comes to you with a simple baseplate without any adjusting devices. The total weight of magnet plus its support is about 1370 lb. It will run on its own power supply at the NSLSII. The beam enters the LBTDPL4800#3 dipole such that the power terminals are on the beam’s left.

1. Receiving Information:

| Received by (MM initials) | SDA |
| --- | --- |
| Date received : (dd-mm-yyyy): | 28-Nov-2011 |
| Checked Magnet Number( LBTDPL4800#3) |  |
| Magnet Engineer (Cherrill Spencer, ext 3474): verifies she has inspected the magnet and authorized the magnetic measurements |  |
| No barcode sticker on this BNL dipole. |  |

Preparation:

2. Power and LCW Connections: Unipolar PS  >=160A , >25 V  required. Cosine-shaped ramp that mimics the ramps produced by an Ethernet-driven controller.

Low pressure LCW system will be needed to cool this magnet in your measurement lab.

3. Magnet Orientation:

A temporary beam direction arrow may be visible on the top or side of the core. Otherwise: note that the beam enters the dipole such that power terminals on the left.  If there is no arrow, mark a thin, small one with a black pen on the side core. Please attach dipole & support securely to table as we may have to do a reproducibility test like we did on the BNL quads.

| Beam-direction arrow in place (initials): | SDA |
| --- | --- |

4. Magnet Alignment:

This magnet has spot faces for 12 Hubbs Spherical Mounts, which carry 1.5” spherical mount retro-reflectors, these SMRs need to be fiducialized to the centerline of the dipole‘s aperture, which should be defined using the core’s horizontal split plane and the width of the 2 poletips. An 8 foot FARO arm should be used to carry out this fiducialization so that the equipment used by SLAC is exactly the same as that used by Brookhaven National Lab.

| Fiducialization of all SMRs complete (Alignment Initials :) | LG |
| --- | --- |

Alignment crew should do the set up on the measurement stand of the dipole and the stretched wire  and use the x,y,z axes of the whole magnet (as defined by SMR measurements) such that the roll angle (angle of horizontal axis) is less than about 2 milliradian (or as small as possible), and pitch and yaw are also minimized.

Fiducialization data on location of the SMRs with respect to the geometric axis when poles are aligned perfectly horizontally should be saved in a WORD or EXCEL document, because this data will be sent to the Survey Group at BNL.

| Alignment completed, Roll Angle = 0.1 millirad (Alignment Initials :) | LG |
| --- | --- |

Your stretched wire system must be set up carefully too. Please make sure its wires are oriented so they are really horizontal relative to the horizontal axis defined by the alignment crew.  We will NOT use a rotating coil on this dipole.

5. Power leads and polarity:

The required polarity for LBTDPL4800#3 is negative horizontal dipole.   Looking downstream the upper- pole should be a NORTH pole as shown in the above figure. Don’t take the position of the terminals in the sketch literally, they will be to the left as one looks downbeam.

Determine the electrical connection polarity (with supply outputting positive current) which produces the required magnetic field polarities.

Place + and  - labels on or near the main power flags for the required polarity.

| Polarity established, power +/- labels applied | SDA |
| --- | --- |

6. Cooling Water Flow Check:

Each dipole has 6 water circuits, 3 per coil, with 1 input and 1 output spigot with appropriate Ys and Ts internal to the quad that divide up the LCW flow into 6 parallel circuits.

Adjust the differential pressure across each water circuit (deltaP) to be about 60 psi.

Total water flow should be about 1.6 gpm. If you cannot adjust deltaP to 60 psi, first use close to 60 psi, measure the flow, and then make total flow 1.6 +/-0.2 gpm by adjusting the valve.

|  | Water flow [gpm] | Check |
| --- | --- | --- |
| Maximum delta P _______ | 1.61 | SDA |
| Adjusted to produce 1.60+- 0.2 gpm: | 1.60 or |  |

7. Thermocouple placement and Thermal Test:

This is a brand new magnet and has never seen any current, so the first time you power it (probably for this thermal test) please bring it up slowly (say, 50 amps increase at each step) and watch the temperature measurements for a minute as each current is reached. At 155 amps I predict an LCW temperature increase of less than 10 degrees C, if the increase is much higher than that then something is wrong and call Cherrill Spencer, ext 3474.

Place 5 thermocouples at these locations:

- any one of the coils – on the incoming LCW – on the outgoing LCW  -elsewhere on bench for ambient temperature measurement – on the steel core near any pole tip.

Arrange for temperatures to be read at least once every 3 minutes for the following thermal test:

After water flow has been checked on each circuit, run JUST the FIRST dipole to 155 A for three hours, reading the thermocouples, magnet current and voltage

| Confirm total LCW flow is 1.60 +/- 0.2 gpm | Flow is 1.61 gpm |
| --- | --- |
| Name of thermal test file and date of test | rtdat.ru1 |
| URL of webpage where datafile stored |  |
| Final LCW temp (in) – Final LCW temp (out) | = 6.3 °C |

If the rise in the LCW temperature is more than 16 degrees C at 155 amps, something is wrong; stop, call Cherrill Spencer at x3474 or her substitute.

8. Initialization:

To be done only once, after successful completion of the water flow check and Thermal Test above. Use the cosine ramp with an equivalent linear rate of 15 A per second; minimum current of 5 A, maximum current of 175 A and pause time of 120 seconds at the minimum and maximum currents. Take the quad from 5 A to 175 A and back 7 times, ending up at 5 A.

| Initialization completed: date , MMG initials: | 12/2/2011, SDA |
| --- | --- |

9a. Standardization parameters and detail: (For use in measurements below)

Ramp type: cosine shape

Minimum standardisation current: 5 amps

Maximum standardisation current : 155 amps

Pause time at minimum or maximum standardization current: 30 seconds

Equivalent linear ramp rate for standardization process: 15 amps/second (increasing or decreasing)

Start at 5 A and ramp up to 155 A and down to 5 A FIVE times, finishing at 5 amps.

Setting to desired current (BDES) : ramp UP to BDES at 5 amps/second

All operating currents should be approached from below; except when I request a series of decreasing currents – these will be approached from a higher current.  When moving from one current to another for the strength or harmonics measurements use 5 amps/second and wait at least 30 seconds after the operating current has been reached before making any magnetic measurements. Put basic standardization details in all datafile comments.

| Confirm that magnet will ramp at 15 A/s (initials:) | SDA |
| --- | --- |

9b. Multipole standardizations BEFORE any magnetic measurements

As this is a brand new magnet it needs some training along the specified standardization cycle before any magnetic measurements are taken. So take it through 6 whole standardizations (6x5=30 up and down cycles) before making any of the measurements listed below.

10. Poletip Field Measurements

The aperture is 48mm tall. The pole is 200.15mm wide, it has no bumps or chamfers on its pole edges.

When it is convenient, maybe when the stretched wire is not in the dipole’s aperture, please use a Hall probe to measure the poletip field  at 141 amps (after a regular standardization procedure) . Please make a correction in the field value for the thickness of the Hall probe.

| Current Actual current | Poletip field : |
| --- | --- |
| 141 A nominal 140.984 A | 0.702 T |
| Date and initials: | 12/2/2011 SDA |

11. Magnetic Measurements of ∫Bdl & harmonics at various main currents and y positions.

Purpose of these measurements is to find the transfer functions (current required to reach a certain integrated strength) especially for certain desired integrated strengths, and to check the homogeneity of the integral strength within the good field region defined by the BROOKHAVEN requirements. Your stretched wire system must be set up carefully. Please make sure its wires are oriented so they are really horizontal relative to the horizontal axis defined by the alignment crew.

For LBTDPL4800#3, the nominal ∫Bdl is -2.29 kGauss-m (predicting a nominal current of 141 A)

The nominal good field region is +-15mm in the y axis, all along z and +-20mm in the x axis, all along z. Within this box the (∫Bdl - ∫Bdl at x=y=0)/ (∫Bdl at x=y=0) must be less than 10x10-4 .

- Integrated Strength Measurements:

Measure the integrated strength ∫Bdl at these currents at the x=0, y=0 center of the poletips:

25, 40, 55, 67, 80, 93, 106, 120, 130, 141, 150, 155 amps

Then back to 25 amps from 155 amps, measure the ∫Bdl at the same set of currents in reverse order. Remember to ramp at 5 amps/second and wait for 30 seconds after reaching the new current before starting a measurement.

| Filename of Int. Strength Measurement | Wiredat.ru2 |
| --- | --- |
| Date of measurement, initials | 12/2/2011 SDA |

I will do a fit of ∫B.dl versus current for the increasing set of currents and see if my predicted current of 141 A for the dipole is close and calculate the actual operating current. Brookhaven have specified the 3 currents they want the field homogeneity measured at: 67, 93 and 141 amps.

- Homogeneity and Harmonics Measurements:

I think BNL will appreciate knowing the multipoles in this dipole as well as the homogeneity of the integral in the specified food field region. So we need to take several ∫Bdl measurements at various x and y positions and at three currents. We need a precise measurement of the harmonics; especially the sextupole and we will use the stretched wire setup to take the necessary data.  The poletip is 200.15mm wide with no bumps or chamfers. I want integrated strength measurements at 17 different x positions at 5 different y planes; and each grid of points to be measured at 61, 84 and 141 amps.

From this data I will fit polynomials to the ∫B.dl versus x plots and I will calculate the multipoles and the homogeneity.

Here are the 17 x positions in centimeters, x=0 is the width center of the poletip:

-3.0, -2.5, -2.0, -1.5, -1.0, -0.75, -0.5, -0.25, 0, 0.25, 0.5, 0.75, 1.0, 1.5, 2.0, 2.5, 3.0 cm

Here are the 5 y positions in centimeters, y=0 is the height center of the gap :

-1.5, -1.0, 0.0, 1.0, 1.5 cm

| Filenames of Harmonics Measurements | Y= 0.0 wiredat.ru3 12/5/2011 Y= 1.5 wiredat.ru4 12/5/2011 Y = 1.0 wiredat.ru5 12/6/2011 Y = -1.0 wiredat.ru6 12/7/2011 |
| --- | --- |
| Date of measurement | Y = -1.5 wiredat.ru7 12/7/2011 |

NOT YET SURE IF HAVE TO DO A MECHANICAL SPLITTING TEST like we did on the BNL quads- will let you know in next couple of days.

12. Saving and Distributing the Magnetic measurement data

Please post all data as it is acquired at

http://www-group.slac.stanford.edu/met/MagMeas/MAGDATA/NSLSII/dipole/ etc ___________________________________________________________________

in a subdirectory with the magnet name.

Please inform Magnet Engineer ( Cherrill Spencer at cherrill@slac.stanford.edu  (ext 3474) when the strength and harmonic measurement data are available for inspection and analysis, and for final approval . If Cherrill is on vacation please send the pertinent thermal, strength and harmonic txt files attached to an e-mail to Raymond Fliller at BNL: rfliller@bnl.gov .

13.             This section is to be completed by Cherrill Spencer, SLAC.

| Data been inspected & analyzed by Magnet Engineer |  |
| --- | --- |
| Based on Magnet Engineer analysis of data can confirm the pre-assigned beamline location (optics-deck name): | Optics beamline location: |
| Dipole is released from Magnetics Group, signed and dated by Cherrill Spencer: |  |
| Remove the LCW from the coils and take the magnet to just outside your cage. Signed & dated by MMG personnel: |  |

14.               This section is to be completed by BNL Beam Physicist (Ray Fliller).

| Checked that integrated strength data is satisfactory and have generated the polynomial function for the controlling database to set the magnet. Nominal operating current is _______________amps. Checked that this is within the capability of the assigned power supply _____________________ | Signature and date: |
| --- | --- |
| Checked that the integrated strength homogeneity requirement is met( initial): |  |
| Magnet measurements accepted for BNL (signed): | Electronically signed: |
| Date measurements accepted (Month-Day-Year): |  |

The SLAC MMG person in charge of the magnetic measurement process is Scott Anderson, his phone number is 650-926-4590, sda@slac.stanford.edu . If Cherrill Spencer is absent from SLAC then Scott Anderson will coordinate all the tasks associated with these measurements.

.When this traveler is completed, attach the original to any other traveler on this magnet and send a photocopy to Cherrill Spencer at MS 12.

This dipole will need to be returned to the Stangenes Company as soon as these measurements are done, so please inform Cherrill Spencer, or in her absence, Lance Thompson at Stangenes, 650-493-0814, lthompson@stangenes.com when the measurements are done.

Further Tasks to be done on this magnet will be recorded on its Fabrication Traveler (owned by Stangenes).