# 7.0	Alignment Proposal

This chapter will describe procedures and methods which, if carried out in a careful and rigorous manor, will yield the aligned position of all LCLS components within their position tolerances. Major geodetic principles governing survey and alignment measurement space are briefly revisited and their relationship to a lattice coordinate system are shown. The chapter then continues with a discussion of the activities involved in a step-by-step sequence from initial lay-out to final alignment.

## 7.1.	LCLS Surveying Reference Frame

Horizontal position differences between the projection of points on the geoid or a best fitting local ellipsoid and those on a local tangential plane are not significant for a project of the size of the LCLS. Hence, it is not necessary to project original observations like angles and distances into the local planar system to arrive at planar rectangular coordinates.

Table t71_a    Curvature Correction

| Distance r [m] | Sphere HS [m] | Spheroid HE [m] |
| --- | --- | --- |
| 20 | 0.00003 | 0.00003 |
| 50 | 0.00020 | 0.00016 |
| 100 | 0.00078 | 0.00063 |
| 1000 | 0.07846 | 0.06257 |

However, in the vertical plane, the curvature of the earth needs to be considered (see fig. f71_a). Since leveling is done with respect to gravity, the reference surface is the geoid. Because of the relatively small area of the LCLS project, one can substitute the non-parametric geoid with a locally best-fitting sphere or spheroid. Table 1 shows the projection differences between a tangential plane, a sphere, and a spheroid as a function of the distance from the coordinate system’s origin. Notice that for distances as short as 20 m the deviation between plane and sphere is already 0.03 mm (see table t71_a).

### 7.1.1	Network Design Philosophy

The global alignment tolerance, the relatively weak links between machine sections, and advances in surveying make it possible to forego the traditional design of a two tiered network hierarchy covering the whole machine. Instead, each machine section can be considered independent, only connected by tunnel networks. A two tiered approach will only become necessary if a remote experiment is added to the project and at that time needs to be connected to the undulator. Omitting a primary network not only removes many constraints for component placement, since fewer lines-of-sight need to be maintained, but also presents a significant reduction in alignment costs.

Traditionally, forced centered “2+1-D” triangulation and trilateration techniques were used to measure the tunnel networks.

Fig. f711_b    Sphere mounted theodolite target

Fig. f711_c Sphere mounted glass and air reflectors

However, a 3-D “free stationing” approach does not require forced centered instrument set-ups, thus eliminating the need for set-up hardware and their systematic error contribution. Removable heavy duty metal tripods, translation stages, CERN sockets and optical plummets are not needed (see fig. f711_a). The network design still must consider other systematic error effects, especially lateral refraction. Another important consideration is the target reference system. The design of such becomes much easier with free stationing since we are dealing only with targets and not with instruments as well. Accordingly, it is proposed to use a design, which is now widely used in high precision metrology. This approach is centered around a 1.5" sphere. Different targets can be incorporated into the sphere in such a way that the position of the target is invariant to any rotation of the sphere. At SLAC, designs have been developed to incorporate theodolite targets (see fig. f711_b), photogrammetric reflective targets as well as glass and air corner cubes (see fig. f711_c) into the sphere. Receptacles for the spheres, which are usually referred to as “nests” or “cups”, have been designed to accommodate different functions. CEBAF has a very suitable design for nests to be grouted into the floor, and designs are available at SLAC for cups tack-welded onto magnets, for mounting cups on wall brackets and for a “centered” removable mounting placed into tooling ball bushings (see fig. f711_d).

Fig. f711_d        Sphere receptacles: floor, component, and wall bracket fixed mount versions, removable

centered version

### 7.1.2	Network Lay-Out

The LCLS network consists of three parts: the linac network, the undulator network and the transport line/experimental area network.

7.1.2.1	Linac Network    The linac network serves a different purpose than the other networks. Since the linac already exists, it does not need to support construction survey and alignment, but rather will only provide local tie-points during the linac straightening (smoothing) procedure.

7.1.2.2	Undulator Network   The undulator network’s overall geometry is dictated by the tunnel geometry, machine lay-out and the fact that the free stationing method requires a greater number of reference points. The geometry should also permit observing each target point from at least three different stations. The reference points can be of two different hierarchical classes. The second order points, or tie points, mainly serve to connect the orientation of free stationed instruments, while the first order points additionally provide the geometric reference during machine installation; they are the equivalent of traditional traverse points or monuments. The following sketch (fig. f7122_a) shows a typical section of the lay-out. A triplet of monuments is always placed in the tunnel cross section containing a corrector magnet. One monument will be placed close to the corrector magnet on the floor, the second one mounted at instrument height to the aisle wall, and the third monument mounted to the back wall also at instrument height.

7.1.2.3	Transport Line/Experimental Area Network   Initially, the project will only include an experimental area directly down-stream of the undulator. This network will be constructed and established similar to the undulator network. The only difference being that each cross-section will have only two monuments. These monuments will be mounted to the walls at instrument height.

### 7.1.3	Alignment Coordinate System

The alignment coordinate system will be a Cartesian right-handed system. The origin is placed at the linac Station100 (in analogy to the SLC coordinate system). There will be no monument at the origin, it is purely a virtual point. The Y-axis assumes the direction of the gravity vector at the origin but with opposite sign. The Z-axis is in the direction of the linac, and the X-axis is perpendicular to both the Y and Z axes. The signs are defined by the right-handed rule (see fig. f713_a).

### 7.1.5	Network Survey

The most efficient instrumentation for the network observations will be a laser tracker (fig. f715_a) /total station combination. We have available two state-of-the-art SMX4000/4500 laser trackers, and in addition a new Total Station optimized for industrial metrology, the Leica TDA5000 (fig. f715_b). While it is superior in angular accuracy to a laser tracker, it is significantly less accurate in distance measurements. Consequently, we will rely on the laser tracker for distance accuracy and on the Total Station for angle accuracy.

The laser tracker will be placed close to the intersection of the diagonals of each reference point quadrilateral (see fig. f715_c). From there, four points in a forward direction and four points in backward direction will be measured. The measurement procedure will include three sets of direction measurements to the same eight points in both front and reverse positions plus one set of distances in both positions. To strengthen the determination, the Total Station will be used to measure additional direction sets. It will be set up at two additional stations at the Z-position of the intersection of the quadrilateral diagonals and laterally shifted by about 0.5 m in both directions. From these stations direction sets to six points each downstream and upstream will be measured. To strengthen the overall straightness of the network, wire offset measurements between the wall monuments will be carried out. To minimize wire motion noise, a 30 m long wire in a 10 m overlap mode will be used. All reference points will also be observed with a standard high precision double-run level procedure. A Zeiss DiNi11 digital level in combination with 2 m invar rods is recommended. To further strengthen the elevation determination, all points will also be observed with the portable HL5 hydrostatic level. First simulations show absolute station standard deviations of less than 80 µm without wires and less than 40 µm with wire measurements.

### 7.1.6	Data Analysis and Data-Flow

To reduce the data from the measurements as described above, special software is required. This type of analysis software is based on the photogrammetric bundle approach. Since a photogrammetric sensor is arbitrarily oriented in space, not only its translational parameters but also its rotational orientation parameters must be treated as unknowns and become part of the solution. With traditional trilateration/ triangulation based analysis software, however, pitch and roll are supposed to be oriented to gravity, and yaw is expressed as a function of translations. Additionally, the traditional software assumes that the instrument is centered on a point to which sufficient measurements have been taken. This analysis approach does not work well with free-stationing, and does not work at all with present generation laser trackers, since they cannot be oriented directly to gravity.

To reduce errors stemming from transcription of data, the data-flow should be automated. The suggested instruments support direct connection to field computers. The fully automated data-flow should extend from field computers through data analysis to data storage.

Measurements with any type of instrument will be guided by software based on rigid procedures running on field data logging computers. The data logging software will also pre-analyze the measurements, and will try to determine and flag possible outliers before the measurement set-up is broken down. This method combined with an automated data-flow will greatly reduce errors and improve measurement consistency and reliability.

## 7.2	Lay-out Description Reference Frame

### 7.2.1	Lattice Coordinate System

The LCLS lattice is designed in a right handed beam following coordinate system, where the positive y-axis is perpendicular to the design plane, the z-axis is pointing in the beam direction and perpendicular to the y-axis, and the x-axis is perpendicular to both the y and z-axes.

### 7.2.2	Tolerance Lists

The relative positioning tolerances x, y, z of the undulator sections, dipoles, quadrupoles and sextupoles are listed in the following table:

###### Table t722_a   LCLS positioning tolerances

|  | x [µm] | y [µm] | r [mr] | x/z [µm/m] |
| --- | --- | --- | --- | --- |
| Straightness of undulator sections | 15 | 30 | 1 | n/a |
| Relative alignment between sections | 50 | 50 | 1 | n/a |
| Global straightness of undulator | 100 | 500 | 2 | 100/10 |
| Linac straightness | n/a | n/a | n/a | 150/15 |
| Quadrupole ab initio | 50 | 50 | n/a | n/a |

### 7.2.3	Relationship between Coordinate Systems

The relationship between the surveying and lattice coordinate systems is given by the building design and machine lay-out parameters. The result is a transformation matrix (rotations and translations).

## 7.3	Fiducializing LCLS Magnets

7.3.1	Traditional Fiducialization

The correct fiducialization of magnets is as important as their correct alignment since an error in either task will effect the particles’ trajectory and cannot be distinguished from each other. Fiducialization can be accomplished either through opto-mechanical and opto-electrical measurements or by using fixtures, which refer to a magnet’s reference features. Detailed descriptions can be found in the literature.

The undulator fiducialization will be an integral part of the assembly and alignment procedure.

## 7.4	LCLS Absolute Positioning

Common to all parts of the machine, free-stationed laser trackers, oriented to at least four neighboring points, are used for the absolute positioning measurements. The tracking capabilities of these instruments will significantly facilitate the control of any alignment operation (moving components into position).

### 7.4.1	Undulator Absolute Positioning

7.4.1.1	Undulator Anchor Hole Layout Survey.   During the anchor layout survey, the anchor hole positions for the girder supports are marked on the piers. It is recommended to fabricate a standard template including all anchor holes at one girder end thus significantly reducing the number of individual lay-out pointings. A Total Station from one free-stationed position can locate and position the template with only two pointings. Before the holes are marked, the location of the template should be checked from a second station. In the sequences of work, the last station can then serve for the n+1 girder as its first station. Specialized software is required to improve the efficiency and reliability of this task.

7.4.1.2	Pre-alignment of Girder Supports and Magnet Movers    The undulator girders will be supported by motorized adjustment systems sitting directly on top of concrete piers. The motorized adjustment systems are based on the Bowden cam shaft design. Two individually controlled cam shaft pairs and two single cam shafts provide five degrees of freedom per girder. The cam shaft design doesn’t compromise the rigidity of the supports and, consequently, doesn’t show a resonance in an undesirable frequency range. This mover system comes in two horizontal slices. The bottom piece consists of a mounting plate, which holds the shafts and stepping motors. The top part is integrated into the girder by mounting the kinematic cams to the girder. The girder is held onto the shafts by gravity.

To accommodate easy installation, the bottom parts of the movers, set to mid range, have to be aligned relative to each other. The required relative position tolerance of these, however, is fairly loose, since the two axes cams are only paired with a single axis cam. On the other hand, to retain as much magnet mover range as possible, the bottom part of the magnet movers should be within 0.5mm of their nominal positions.

To facilitate placing a pedestal such that its top is within 0.5 mm of its nominal position, a widely used method can be used. Here, the base plate of the bottom part of a mover is mounted to the pier by four standoff screws, which are grouted/epoxied into the concrete. The vertical/horizontal pre-alignment of the base plate is accomplished by the following sequence of steps: After the four bolts are epoxied into the concrete, a nut with a washer on top is screwed onto each bolt. These nuts are set to their nominal heights by a simple level operation. Next the base plate is set on the nuts, and a set of washers and nuts is then screwed on the bolts to fasten it down. However, the top nuts remain only hand tight at this point. Next, the elevation and tilts of the base plate are set by adjusting the position of the lower nuts, and subsequently checked with a level with respect to local benchmarks. Then a total station with a “free station Bundle” software package is used to determine the horizontal offset and to simultaneously double-check the vertical offset of the base plate from its nominal position. Finally, the base plate is moved into horizontal alignment using a clamp-on adjustment fixture (push - push screw arrangement), and the nuts are tightened to the prescribed torque. To vibrationally stiffen the set-up, the space between the pier and the base plate should be filled with non-shrinking grout after the alignment has been confirmed.

7.4.1.3	Quality Control Survey    Once the above step is completed, the mover positions will be mapped. If the positional residuals exceed the tolerance, a second iteration can be “jump started” by using the quality control map to quantify the position corrections, which need to be applied. Should a second iteration be necessitated, a new quality control survey is required after completion of the alignment process.

### 7.4.2	Transport Line and Experimental Area Absolute Positioning

The absolute positioning of these components will follow the same procedures as described above.

## 7.5	Relative Alignment

### 7.5.1	Internal Undulator Alignment

7.5.1.1	Introduction   The undulator is in the strictest sense a precision instrument composed of 6400 pole/magnet pairs in two equal and opposing rows. The centroid of the field produced by these two opposing rows of pole/magnet pairs must be straight within approximately 50 µm over any 1.92-m section and within ± 500 µm vertically or ± 100 µm horizontally over the entire 100 m. This means that the field centroid must stay within an imaginary cylinder, of 50 µm radius. This 50-µm-radius cylinder can bend, but not more than 50 µm in 2 m and it must always stay within the bounds of a perfectly straight 500 µm by 100 µm elliptical tube. In operation the undulator will receive feedback from the x-ray beam to aid alignment, however non-field-based alignment is necessary to initiate and maintain alignment. An essential task is to relate the magnetic field centerline to non-field-based alignment fiducials. Equally important is the ability to build and maintain a straight field centerline between fiducials.

7.5.1.2	Fabrication and Assembly Process of Undulator Magnets   Since the fabrication and assembly process is intrinsically correlated to a successful fiducialization of the undulator magnets, the description of this step is included here.

Alignment of the undulator starts with the precision fabrication of the poles (Vanadium Permendur) and the magnets (NdFeB). The magnets are roughly 50 mm square by 5 mm thick and the poles are slightly smaller. The magnets and poles are shown in Figure f7512_a. The taper on the pole results in a quadrupole field as discussed in the section on magnetics. These are the only two forms necessary because the magnet is symmetric and the pole can be flipped horizontally and vertically to achieve the necessary field arrangement.

Both NdFeB and Vanadium Permendur are brittle materials that are expected to grind like a friable material, such as Alumina. Hard ceramics, such as Silicon Nitride or Zirconia are ground to sub-micrometer tolerances for applications in diesel fuel injectors and the magnets and poles could be ground by similar methods. Metal bonded diamond grinding wheels may be used for such precise grinding. The form of the wheel is maintained by EDM truing in situ. However, because of the variability of the magnetic properties of the materials there is no reason to hold exceptionally tight overall dimensional tolerances. Flatness and perpendicularity to adjacent surfaces is important but this is readily achieved.

Physical tolerances are important to the assembly of a straight magnetic field centerline and uncertainties in the physical dimensions of the pole and magnet are minimized via the bonding of a pole to a magnet. A very precise jig is made to kinematically hold the pole and magnet so that the pole face (beam side) is precisely located relative to the bottom and back (opposite the beam) sides of the magnet. The bottom and back sides of the magnet will be aligned by mounting the pair on precisely machined rails of a strongback.

Kinematic mounting, i.e. 6 points of contact, is provided for both the pole and the magnet within the jig (see fig. f7512_b). There are actually 7 points that the pole could contact. The 7th point, which can be seen on the lower left side of Figure 2 AA, allows the pole to be flipped vertically so the flat surface is on the bottom. Only 3 points of contact are shown for the magnet in view BB. The other 3 points are formed by shims on the pole. The shims on the pole ensure a uniform adhesive gap thickness between the pole and magnet. A spring is shown holding the magnet and pole against their kinematic contacts, whereas in reality individual forces would be applied directly opposite each contact to prevent distortion. The important point to note is that the beam-side-pole-surface is precisely related to the back and bottom side of all 6400 magnet/pole pairs. The same relationship will hold for all 6400 magnet/pole pairs since they will use the same or identical bonding jigs.

The magnet/pole pairs can now be fiducialized by scanning the shape of the poles on a high accuracy coordinate measuring machine (CMM). Alternatively, a micro-size Hall probe in place of a standard probe on the CMM could be used to map the field. By correlating the fields of a F/D pair, a centerline can be derived. Magnet/pole pairs can be sorted to minimize the impact of magnetic properties on the field centerline within a given 1-m-module or similarly minimize the effects of one module relative to the next.

Strongbacks or girders are roughly 2 meter aluminum extrusions with a host of critical features. The extrusions are made of 6061, with several modest tolerance features such as coolant fittings, threaded bolt holes, tooling ball holders, and adhesive channels machined in. After rough machining, the extrusion is heat treated and tensioned to stabilize it in the same way aluminum tooling plate is stabilized after rolling. The strongbacks will also be allowed to age for a few months while mounted on the kinematic mounts. Six small rails run the length of the strongback.  These rails form the kinematic contacts for the bottom and backside of the magnets. The space between the rails is an adhesive gap of about 75 µm thickness. Excess material is left on the rails when they are rough machined, so that some material will need to be removed after the stabilization process. Coolant fittings and bolts would be installed prior to final machining. Coolant could be circulated through the strongback while it was being machined and it would be held on the machine bed using the kinematic mounts. The manufacturer of a SIP 5-axis mill guarantees 2 µm accuracy over 1 meter. Final machining of tooling. ball holders would also be done at this time, to assure precise alignment relative to the rails. Tooling balls are provided as an optical alignment reference, because optical techniques have always been an issue for similar instruments. The concept of the strongback is shown in Fig. f7512_c, without tooling ball holders, bolt holes, or kinematic mounts

The rails determine three degrees of freedom for each magnet, but pitch, yaw, and Z are determined by contact of one magnet/pole to the next. There is a gap along the Z axis of approximately 50 µm between each magnet/pole pair. This gap is set by three shims and is filled with a low shrinkage adhesive. Buildup of errors in the z-axis can be monitored with the CMM and shims can be altered to minimize accumulated errors. An auto-collimator will be used to monitor roll and pitch when setting the clamping force (three clamps, each acting against one shim) in order to prevent unwanted pitch and yaw deflections while the adhesive is curing.

Error buildup from various sources is non-systematic, and thus should be added in quadrature. The errors in the physical alignment of a 1-m section are summarized in Table t7512_a for two cases. The first case is based on manufacturer specifications of the best available equipment and should be achievable after modest R&D effort. The second case is based on results from presently available manufacturing services.

Physical alignment of the poles of a 1.92 m section can be directly measured with an appropriate CMM or derived from field mapping data. Mapping data would require scanning with a micro-sized Hall probe attached and spatially calibrated to a CMM. This is the start of an iterative process to further reduce the nonlinearity of the magnetic field centerline. Magnetic shims can be added to the back side of the Vanadium Permendur poles to move the field centerline in both X and Y, i.e. the transverse directions. The linearity of the field centerline could approach the accuracy of the CMM and the Hall sensor.

7.5.1.3	Fiducialization of 1 m long Undulator Sections   The alignment of the undulator will rely on stretched wires simulating a straight reference line.  The fiducialization process will position absolute Wire Position Monitors (WPMs) on each section such that final alignment is achieved by placing a section at its designated position and adjusting its position until the WPMs read zero.

Steps required relating the magnetic field centerline, BPMs, and XPMs to the stretched wire reference can best be understood with a simplified preview of how the system works. Imagine that one would like to align a series of spools, each with a hole through the middle. Imagine that the hole fits perfectly over a wire that can be stretched perfectly straight. Aligning the spools would be extremely easy in this case, but what if the hole could not be put through the middle (see fig. f7513_a part (a)). If two holes could be put through the spool, with the exact same offset, then the spools could be aligned on two perfectly straight stretched wires (see fig. f7513_a part (b)). Now, imagine that the wires sag over the 100 meters that they are stretched. Fortunately, the amount of sag in the wires can be predicted very accurately, so the centerline of the spools can still be aligned by knowing where each spool will be placed along the 100 m wires. The y-offset, or wire sag, is determined for each end of each spool and the alignment holes are put in differently for each spool (same x-offset but different y-offset). Now the centerline of the spools can be aligned by sliding them onto the wires in the proper order. Spools must now be replaced by 1-m-sections of undulator, and a non-contact sensor will have to substitute for the small holes that thread over the wire.

Wire sag, or catenary deflection, can be calculated, but the accuracy of the calculations is limited to the accuracy of the input data. The form of the catenary is

where y is the vertical displacement, T is the horizontal tension on the wire, w is the specific weight or weight per unit length, and z is the displacement from the middle (assuming the ends are at the same elevation). For an extremely flat catenary, the difference between the hyperbolic cosine expression and a parabolic expression is less than 0.5 µm.

BeCu is a good candidate material because of its high strength to weight ratio and the fact that it can be manufactured with excellent uniformity. Based on the bulk properties of BeCu, w is approximately 2.58 x 10-3 N/m and T is approximately 25 N at 70% of yield. If the wire is 102 m long (to include the beam transport approaching the undulator) then the deflection of the wire is 134.2 mm at the middle (x=51 m). To know the shape of the wire at any x position to within 5 µm, the tension and specific weight would have to be known to better than 1 part in 38,000 (• 134/.005). However, if the deflection of the wire is measured at several locations, then the shape of the wire can be predicted over the entire length to better than 5 µm and to within 1 µm over any 2 to 10 m section.

Measuring the catenary deflection of the wire is a task ideally suited for a hydro-level. A hydro-level is in essence a pool of fluid extending the length and width of the undulator and thus defining a horizontal X-Z plane. To measure the y-offset of the wire relative to the fluid requires two different types of non-contact sensors. The surface of the fluid could be sensed with capacitance gauges or laser proximity gauges. The wire position can also be sensed with a variety of non-contact WPMs. A WPM and a fluid sensor would be attached to a tooling bar so the separation between the two sensors could be adjusted to specific values associated with z-axis positions. A minimum of 3 WPM/fluid offset sensors would be used along each of the two wires (more would actually be used to improve accuracy).

The y-position of the WPMs and fluid sensor offset tooling (independently controlled at each of the assemblies) would be maintained using a control loop that keeps the fluid sensor in a constant position relative to the fluid. A separate control loop would maintain the height of the wire at its ends. The fact that the undulator is built on a 5 mrad sloped plane poses no serious difficulty and can be accommodated. Finally, using the WPM at the Z2 position (see fig. f7513_b) the tension on the wire can be controlled to maintain a predefined catenary.

The curvature of the wire is defined by the separation of the end points, specific weight, and tension of the wire. The curvature of the 102-m-wire, over any given section, can be simulated with a much shorter wire by using the same tension and by holding the ends of the short wire at heights corresponding to the section of 102-m-wire. This is a very useful characteristic in positioning the WPMs on each of the 1-m undulator sections.

Horizontal separation of the two reference wires is achieved with micrometer accuracy using specially ground rollers. The rollers would be approximately 2 cm in diameter and 20 cm long. They would contain 3 grooves to match the radius of the reference wires and an additional wire, which could be used to calibrate the BPMs. Spherical receptacles like those shown in fig. f711_d will be added to each side of the roller to facilitate determining the spatial position of the roller on a CMM by probing inserted spheres. Knowing the spatial position of the roller in combination with measuring the position of the grooves on the roller relative to the roller’s axis defines the spatial wire position. Consequently, the wire position would be known relative to the CMM axes and reference position. The rollers need to be able to move up and down on stages with linear encoders. The CMM will provide the necessary input to ensure that the rollers are parallel and that they both travel vertically. A special WPM can be built as a substitute for the usual CMM probe to confirm the position and the catenary of the wires.

Each 1-m section of the undulator is built on the CMM with its magnetic field centerline coincident with two small sections of undulator magnets, which are rigidly attached and aligned to the CMM (see fig. f7513_d). Mechanically probing the pole surfaces or measuring the field of a pole pair with the previously introduced special Hall probe attached to a CMM allows the magnetic centerline of the two reference sections to be aligned to the axes of the CMM. Fig. f7513_e shows a top view of the CMM layout.

With the 1-m section aligned to the reference magnets on the CMM, the reference wires are moved up or down so that the curvature of the wire matches the curvature and position of the 102-m reference wires for this particular 1-m section. WPMs can now be attached and moved into place.

The WPMs are simply a U-shaped bracket with a light emitting diode on one side and a split silicon photo diode on the other side. The wire runs between the two diodes and casts a shadow on the silicon photo diode. If the WPM is off-axis relative to the wire, as shown on the left of Fig. f7513_f, then there is a large differential output. When the diode is centered about the wire, the output of the two diodes is the same and a differential amplifier indicates a null. A small wire diameter promotes better resolution without hurting dynamic range because it is clear which side of the detector the shadow is on.

Each WPM is composed of a pair of sensors like the one shown in Fig. f7513_f. One sensor measures displacement from the wire in the X direction (shown) and the other sensor measures displacement in the Y direction (see Fig. f7513_g). The two sensors are contained in a package that can be attached to the strongback of each 1 m long undulator magnet over a vertical range of approximately 160 mm. The sensors need to be moved until they are centered on the CMM reference wires. Fine adjustment of the sensor could be achieved using flexure with a 1/4-28 screw threaded through one side and a #6-32 screw fixed in the opposite side and threaded into the larger screw. A 5 degree turn of the larger screw pulls the flexure together by about 1.5 µm and translates the sensor about 1 µm.

The WPM package would allow the wires to be moved into the sensor and would provide a means of attaching a protective tube. The protective tube would run from WPM to WPM and prevent air disturbances. The connection of the pipe would be physically isolated from the sensor.

Every fourth strongback should be 25 cm longer to allow a BPM, an XPM, and a steering quadrupole to be attached to it. The BPM can be aligned to the centerline of the magnets using the third wire on the precision ground rollers. The position of this wire is known relative to the magnetic field centerline and a current pulse through the wire would simulate the electron beam. The XPM and steering quadrupole could be aligned using physical datums.

7.5.1.4	Assembling the 1.92-m sections into a 100-m undulator Each section is fabricated to produce a straight magnetic field. A minimum of 3 WPMs are attached to each section to provide the necessary information to align the section in X, Y, xy andz. The 4th WPM provides some redundancy and thus improves accuracy and reliability. The sections are assembled into a continuous undulator along a path defined by surveying techniques. Relative alignment of ± 2 mm to the grid monuments (see chapter 7.1.2.2) will be sufficient at this point.

The end points for the wires will be established by surveying. Precision ground rollers -the same as those that hold the wires on the CMM- are put in place to hold the wires at the end points. The wires are now strung through all the WPMs and the protective tubes are put in place. The wire is now pulled to the anticipated tension and a visual check is made to assure that the wire is not in contact with any of the WPMs. The first WPM/fluid sensor, which is very close to the first roller, is moved to center it on the wire. Now, the fluid level in the hydro-level is adjusted so that the fluid sensor is at the proper offset. Next, the last WPM/fluid sensor would be moved to put the fluid sensor in the correct position relative to the fluid. This action would ideally center the WPM on the wire, but if not the wire height should be adjusted so the anticipated catenary is generated. Finally the tension on the wire should be adjusted so that the predicted offset between the wire and fluid is satisfied at the middle. Changing the tension in the wire should cause a minimal change in the wire height at the ends, although some iterations may be necessary.

By following the procedure in the last paragraph one will establish the position and shape of the reference wires. Now, with each 1-m section in its proper z-axis location, the WPMs will provide the necessary feedback to align the magnetic field relative to the wires.

There are about twenty 2-m-sections of the undulator and each section must occupy a very specific location along the z-axis.  The z-axis spacing is important for the phase matching of the beam but also for proper positioning of the WPMs.  The actual length of a “2-m-section” may be closer to 1.92 meter and, as discussed earlier, there will be gaps to accommodate the BPMs, XPMs and steering magnets.  The actual z-axis position of every section needs to be planned to within 0.2 mm because the slope of the reference wires, , is significant towards the ends and an error in z translates directly into an error in y-position.  As an example, however, consider a system where every section is 1.000 m long the reference wires are stretched 102 meters.  The undulator is centered between the ends of the wires. If we choose the y-offset to be zero at the middle of the undulator and if the WPMs are set 50 mm from the ends of each section, then Table 2 shows the offsets for the first 10 sections, and several toward the middle.

Table t7514_a.  WPM offsets in the y-direction

| WPM # | z-position | y-offset (mm) | Wire slope |
| --- | --- | --- | --- |
| 1A | 1.05 | 128.7421 | -0.00515 |
| 1B | 1.95 | 124.1446 | -0.00506 |
| 2A | 2.05 | 123.6389 | -0.00505 |
| 2B | 2.95 | 119.1342 | -0.00496 |
| 3A | 3.05 | 118.6388 | -0.00495 |
| 3B | 3.95 | 114.2270 | -0.00486 |
| 4A | 4.05 | 113.7420 | -0.00485 |
| 4B | 4.95 | 109.4231 | -0.00475 |
| 5A | 5.05 | 108.9484 | -0.00474 |
| 5B | 5.95 | 104.7223 | -0.00465 |
| 6A | 6.05 | 104.2579 | -0.00464 |
| 6B | 6.95 | 100.1248 | -0.00455 |
| 7A | 7.05 | 99.6707 | -0.00454 |
| 7B | 7.95 | 95.6304 | -0.00444 |
| 8A | 8.05 | 95.1866 | -0.00443 |
| 8B | 8.95 | 91.2392 | -0.00434 |
| 9A | 9.05 | 90.8058 | -0.00433 |
| 9B | 9.95 | 86.9513 | -0.00424 |
| 10A | 10.05 | 86.5282 | -0.00423 |
| 10B | 10.95 | 82.7665 | -0.00413 |
| 48A | 48.05 | 0.4490 | -0.00030 |
| 48B | 48.95 | 0.2168 | -0.00021 |
| 49A | 49.05 | 0.1962 | -0.00020 |
| 49B | 49.95 | 0.0569 | -0.00011 |
| 50A | 50.05 | 0.0466 | -0.00010 |
| 50B | 50.95 | 0.0001 | -0.00001 |
| 51A | 51.05 | 0.0001 | 0.00001 |
| 51B | 51.95 | 0.0466 | 0.00010 |
| 52A | 52.05 | 0.0569 | 0.00011 |
| 52B | 52.95 | 0.1962 | 0.00020 |

Each 1-m-section has four WPMs, two on the left and the right, as viewed looking down the axis of the machine (see fig. 7513_b).  The y-offset on the left and right are the same, so they will be discussed as if there are only two per section.  The WPMs can be designated as A and B so the first section would have WPMs 1A and 1B, etc.  The slope of the wire is listed in Table t7513_a to illustrate the importance of accurately locating each section in the z-direction, especially toward the ends.  At the ends, a 1 mm error in z would produce a 5 µm error in y, which is unacceptable.

The problem with aligning the undulator two-fold: first the magnetic centerline is difficult to detect, and second, the alignment reference is off axis. Errors can enter into the alignment through a host of measurements needed to relate the magnetic field line to the reference line. Table t7513_b summarizes the sources of errors and gives estimates of the magnitude of the error, at two levels of confidence. The estimates in Table 3 are made on the basis of reported resolution, CMM error mapping, and a fair amount of engineering judgment.  None of these error levels will be easily achieved and considerable effort should be expended in proving that these error levels can be achieved consistently.  An important point is that there are no dominant sources of errors and thus many small errors can be added in quadrature, with minimal contribution to the error budget.

Table t7513_b. Error sources in the magnetic field alignment.

| Source | Moderate Confidence | High Confidence |
| --- | --- | --- |
| Hall probe sensitivity (resolution) | 0.3 µm | 0.9 µm |
| Hall probe relation to CMM | 0.5 | 2.0 |
| CMM Accuracy | 1.0 | 1.0 |
| Ability to shim field | 1.0 | 2.0 |
| (dy/dz)∆z where ∆z is the z error on the CMM | 0.1 | 0.1 |
| (dy/dz)∆z where ∆z is the operational error | 0.5 | 1.0 |
| WPM sensitivity | 0.5 | 0.5 |
| WPM amplifier drift | 0.5 | 1.0 |
| WPM spatial positioner resolution | 0.5 | 1.0 |
| Hydro-level fluid | 1.0 | 2.0 |
| Fluid level sensor | 1.0 | 2.0 |
| WPM/fluid sensor combination | 1.0 | 2.0 |
| Bending of the strongback | 0.5 | 1.0 |
| Bending of the girder | 1.0 | 3.0 |
| Wire catenary calculation | 1.0 | 3.0 |
| Root sum of squares | 2.93 | 6.63 |

### 7.5.2	Linac Smoothing

7.5.2.1	Purpose of linac smoothing   To generate an optimal beam for injection into the undulator, the present local straightness of the linac is not sufficient. To achieve the desired beam parameters, the straightness quality needs to be mapped, and where necessary mechanically adjusted. In particular, the straightness of individual linac structures, the straightness alignment of structures on a girder, and of the relative alignment of the sections on either side of a quadrupole with respect to each other and with respect to the all need to be mapped.

7.5.2.2	Linac straightness measurement procedure   Because of the required resolution, reliability and the large amount of work (about 1 km of beam line), the task is best performed with a system which does not require an operator to point, and adjust micrometers. It also should allow on-line data logging. It is therefore proposed to use a laser system developed by Hamar. The instrument generates two laser light planes by bouncing a laser beam off rotating mirrors. The two light planes are truly perpendicular to each other. The flatness or wobble-induced error of each light plane is specified as 5 µrad, which is well below the straightness specification at maximum distance. The light source would be set-up at about the middle between two quadrupoles offset horizontally and vertically such that the light planes clear all beam line components. This set-up vs. a set-up at an endpoint reduces the length of the line-of-sight to about 5 m thus greatly lessening the effects of potential refraction and air turbulence on the light beam. After bucking in the light planes both horizontally and vertically to two points on the measured object, intermediate offsets between e.g. the accelerator structure and the light planes are measured with a photo-sensitive-detector (PSD) attached to an offset arm. The detector is linked to an interface box by a cable, which can be as long as 15 m. The interface box provides a serial link to a data logger. To measure the offsets, the offset arm is held against the accelerator structure sequentially in both planes. To determine the perpendicular offset, the alignment technician will arc the arm. While the arm is being arced, the light position is continuously read-out and stored. Software will then determine the perpendicular offset by finding the smallest read-out value. Since the PSD measurement range is limited to about 8 mm, the arm will be adjustable in length. To avoid errors due to the adjustability, the adjustment length will be monitored by an electronic dial gauge, which also reports its reading to the data logging software. The total straightness measurement error budget is expected to be below 75 µm.

The relative alignment of a linac quadrupole in relation to its adjacent accelerator sections will be determined analogously. However, since these quadrupoles are not fiducialized and also don’t have any precision reference surfaces, the offset will be measured to their BPMs instead. Each BPM has a cylindrical body, which is inserted between the poles with a very close fit, and protrudes from the poles on either side of the magnet. The BPM is expected to reference the magnet’s axis to about 100 µm. The adjustment range of the offset arm will be sufficiently dimensioned to allow the same arm to measure both BPMs and accelerator structure offsets.

### 7.5.3	Relative Alignment of Transport Line and Experimental Area Components

The position tolerances of these components will be achieved during the absolute alignment step. A relative alignment is not required.

## 7.6	Undulator Monitoring System

In the vertical dimension a hydrostatic leveling system cloned after ESRF’s HLS will accurately monitor relative and global vertical position changes. Four sensors per girder are required. To eliminate temperature effects on the hydrostatic leveling results, any vertical deviations from a plane of the water runs must be avoided. This condition is guaranteed using the “half-filled” pipe approach. Assuming 4 sensors per girder, a total of about 88 sensors will be required.

Since there is no natural absolute reference in the horizontal plane, some kind of artificial local reference needs to be created as is done by stretched wires. In principle, the wire alignment system established during the internal undulator alignment could be used to monitor long term motion of the undulator girders. However, it is not clear at this point of the design, whether the optical sensors can withstand the radiation emanating from the electron beam. It is therefore suggested to include into the machine design the substitution of the optical sensors with low frequency inductive sensors.

# Endnotes

Notes and Bibliography

The Geoid is the reference surface described by gravity; it is the equipotential surface at mean sea level that is everywhere normal to the gravity vector. Although it is a more regular figure than the earth’s surface, it is still irregular due to local mass anomalies that cause departures of up to 150m from the reference ellipsoid. As a result, the geoid is nonsymmetric and its mathematical description nonparametric, rendering it unsuitable as a reference surface for calculations. It is, however, the surface on which most survey measurements are made as the majority of survey instruments are set-up with respect to gravity.

The reference ellipsoid is the regular figure that most closely approximates the shape of the earth, and is therefore widely used in astronomy and geodesy to model the earth. Being a regular mathematical figure, it is the surface on which calculations can be made.

For more information see also: Ruland, R.: Magnet Support and Alignment, in: H. Winick, Editor, Synchrotron Radiation Sources - A Primer, pp. 274 - 304

Forced centering refers to a specific instrument mount. This type of mounting system, whether vendor specific or independent, allows the exchange of instruments on a station without loosing the measurement point, i.e. all instruments are by mechanical “force” set up in exactly the same position. However, experience has shown that even the best of these forced centering systems has a  of about 50-100 µm. Unfortunately, the forced centering system contributed error is not random. Since a whole set of measurements is usually completed from a slightly offset position, this error behaves mostly systematically. No efficient method is known to determine the offset vector. These errors, vertical refraction, and lateral refraction are the biggest contributors to the systematic error budget in surveying engineering.

2+1 -D refers to the fact that because of mechanical problems in the forced-centering hardware, three-dimensional networks were usually split into separate horizontal (2-D) and vertical (1-D) networks. Both networks were established, measured and analyzed separately.

Rather than setting up the instrument over a known point, the instrument’s position is flexible and chosen only following considerations of geometry, line of sight and convenience. To determine the instrument position, at least three points, whose coordinates are already known or are part of a network solution, need to be included in the measurements.

Lateral refraction is caused by horizontal stationary temperature gradients. In a tunnel environment, the tunnel wall is often warmer than the air. This creates vertical stable temperature layers with gradients of only a few hundredth of a degree Celsius per meter. If one runs a traverse close to a tunnel wall on one side only, the systematic accumulation of the effect can be significant. E.g. during the construction of the channel tunnel, a control measurement using gyro theodolites revealed that after about 4 km they had already veered about 0.5 m off the design trajectory.

The " character indicates inches; 1 in = 2.54 cm, hence the diameter of the 1.5" sphere is equivalent to 3.81 cm.

While a wire offset measurement arm with the desired resolution is not available off the shelf, wire systems with even higher resolutions have been built successfully in the past both at LLNL and SLAC

Ruland, R., Setting Reference Targets, in Proceedings of the CERN Accelerator School on “Magnetic Measurements and Alignment”, Capri, 1997, in print.

Precision Instruments in Minneapolis specializes in grinding to such fine tolerances

Typical Hall probes are roughly 1 mm on a side and average the field over that characteristic dimension. With such a probe it might be hard to achieve 1 µm resolution. However, Bell Labs has recently reported on “scanning Hall probe microscopy” with 0.1 gauss sensitivity and 0.3 µm spatial resolution. Also, the Hall probe needs to be calibrated with respect to coordinate axes of the CMM. To facilitate this, the probe could be built into a sphere that could be rotated about the X and Y axes, so the axes of the Hall probe could be determined in relation to the sphere center. This would allow the Hall probe measurements to be related to the physical measurements of the CMM.

E.g. the digital Elcomat 2000 by Möller-Wedel, Germany, with 0.14 arcsec resolution would be well suited.

The Leitz PMM CMM, operated by the SLAC Metrology Department, meets the accuracy requirements. It has a certified accuracy of 0.8 µm over the required 1x0.3x0.3 m volume in clean room conditions and with 0.5( C temperature control. However, the present housing does not fulfill the here necessary cleanroom conditions.

D. Martin, D. Roux, Real Time Altimetric Control by a Hydrostatic Leveling System, in: Proc. of the Second International Workshop on Accelerator Alignment, DESY Hamburg, 1990, pp.171-180.

Hamar Lasersystems, Model L-723 Triple Scan® Laser with Model A-517 Scan Target.

# Text boxes

#### Fig. F7122_a    Undulator network lay-out

#### Fig. f713_a    Coordinate system definition

#### Fig. f715_a    SMX  Laser Tracker

#### Fig. f715_b   Leica TDA5000 Total Station

#### Fig. f715_c   Observation plan

Fig. f7512_a.  The magnets and poles are all the same but the pole will assume different orientations.

Fig. f7512_b  A jig is used to achieve precise orientation of the field shaping pole surfaces to the physical alignment surfaces of the magnet.

Fig. f7512_c  The strongback will be handled by its mounting surfaces (not shown) to prevent distortion.

#### Table t7512_a. 1-m section manufacturing and assembly tolerance budget

| Source | Modest confidence [µm] | High confidence [µm] |
| --- | --- | --- |
| Pole to magnet back surface | 1 | 2 |
| Pole perpendicularity | 1 | 2 |
| Magnet perpendicularity | 1 | 2 |
| Rail straightness | 2 | 4 |
| Rail distortion (handling) | 1 | 3 |
| Attachment | 2 | 5 |
| Sum in quadrature | 3.5 | 7.9 |

Fig. f7513_a  (a) Weightless spools aligned along an ideal stretched wire.  (b) If the wire cannot go through the center of the spools, then two wires are needed and if the wires sag, then they must go through each spool with a different y-offset.

Fig. f7513_b  Calibrated offset tooling is used to maintain the ends of the wires at a prescribed height above the level and to assure a specific catenary by controlling the tension in the wire.

Fig. f7513_c  A precision ground shaft allows the wires to be held with a known relationship to the CMM.

Fig. f7513_d  Two small groups of undulator poles are attached to the CMM so that all the 1-m sections are aligned relative to these references before the WPMs are attached and the BPMs are calibrated.

Fig. f7513_e.  CMM with 2-m undulator section.

Fig. f7513_f  Photo diode type of WPM with micrometer adjustment mechanism.

Fig. f7513_g  Two sensor arrangement

Fig. f7514_a  1-m-section with WPMs.

Fig.f7514_b  First four sections of the undulator aligned to the reference wire.

Fig. f71_a Effect of earth curvature

Fig. f711_a    Forced Centered Set-up at SLAC

#### Fig. f7521_c   Relative straightness alignment of accelerator structures and associated quadrupole

Fig. f7521_b   Straightness of accelerator structures on a girder

Fig. f7521_a   Straightness of individual accelerator structures