Magnet Measurement Test Plan

Fill

Logistic Info

| Person Responsible | Domenico Dell’Orco |
| --- | --- |
| Phone # | cell. 510-219-0378 |
| Start Date | 8/12/2019 |
| Finish Date | 8/16/2019 |
| Acct # for alignment | 14427-A8001 (Septum project) |

Magnet Designation and Type

| Project Name | BTS C8H magnet (2019) |
| --- | --- |
| Magnet Type | Horizontal Corrector |
| Magnet Name/Number | C8H (2019) |
| Serial # / PC # | 01 |
| Trim Coils | No | Trim Types |  |
| Magnet gap or diameter (mm) | 30 mm (1.181”) |

Alignment

| Coarse (1.0 mm)/Precision (0.05 mm) |
| --- |
| Has the magnet been fiducialized | Yes, use same coordinate system as in SPEAR HV correctors. |
| If it has not been fiducialized, what is origin point of the magnet, what is its coordinate system and what surfaces should be used for alignment? |  |
| + Z direction of magnet (Please provide drawing) | Beam direction (Z+) is Stenciled |

Special Alignment Tolerances

| X | Use same as for the SPEAR3 HV Correctors |
| --- | --- |
| Y |  |
| Pitch |  |
| Yaw |  |
| Roll |  |

Water and Power.

| Water Pressure | Air cooled | PSI |
| --- | --- | --- |
| Water Flow |  | Gpm |
| Maximum Current Main | 7.5 A | Amps |
| Max Current Trims | N/A | Amps |

Polarity

| Main Coil Polarity. (Nearest North pole to Horizontal look in +z direction) | Power leads are marked. Positive polarity: B field coming out of the two coils vertically points upward. |
| --- | --- |
| 1st Trim Polarity (Nearest North pole to Horizontal look in +z direction) | N/A |
| 2nd Trim Polarity (Nearest North pole to Horizontal look in +z direction) | N/A |

Temperature Tests

| Temperature and Resistance Test | Yes |
| --- | --- |
| Areas to be measured | Check coil temperature far from cooling plates at maximum current (72 W total at 6.5 A, R= 1.7 Ohm @ 40 C). |
| Currents (Amps) | 7.5 A |
| Check water flow | N/A |

Standardization

| Magnet warm up. ? time (mins) @ ? current (Amps) | Yes, see below. |
| --- | --- |
| Conditioning Standardization Cycles Performed before measurements: | Yes |
| Number of Stdz cycles | 3 |
| Ramp Rate (A/s) | 7.5 A/min | Minimum Current (A) | -7.5A (neg) |
| Settle Time (sec) | 120 s | Maximum Current (A) | +7.5A (pos) |
| Ramp Type (Linear or Three Linear or Cosine | Linear |

WIRE Measurements

| Integrated Strength | Main | Yes | Trims | N/A |
| --- | --- | --- | --- | --- |
| From x=-20 mm to x=20 mm in steps of 2 mm, Y=0 | Current: From -7.5 A to +7.5 A in steps of 0.5 A |  |
| Harmonics | Main | Yes | Trims | N/A |

Harmonic Tolerances (as % of main, Main =100 %)

| Dipole n = 1 | Quadrupole n = 2 | Sextupole n = 3 | Octupole n =4 |
| --- | --- | --- | --- |

Field Map

| Type of map? (Transverse, axial, 3D, Locate Center) |  |
| --- | --- |
| Dimensions ( X, Y, Z) |  |
| Spacing between step (X, Y, Z) |  |

Notes and Additional Tests: