NLCTA Dipole Magnet

FIDUCIALIZATION REPORT

Inspector: 	Keith Caban

Engineer: 	Jim Spencer

Date: 		09/13/05

Charge No:	12-1410-1

Mag. S/N:	A

Fiducialization Coordinate System Set Up:

Magnet is set on 2 granite parallels as displayed in the photo below.

SIDE VIEW

Z zero is set by the mid-plane of these 2 surfaces.

(Distance width = 6.0085”)

Back/Exit View                                             ZOOM Back/Exit View

X & Y zero shown, with X offset -1.000 see (SK-235-978-30)

off of “True Magnet Center”.

---- Dashed Green is Mechanical Centerline

3-D VIEW

Tooling Ball Fiducials

- using ½” Diameter Tooling Balls w/ 1” Extension

Coordinates of Solid Green (Beam)

| TB | FORM | DIAMETER | X | Y | Z |
| --- | --- | --- | --- | --- | --- |
| TB 1 | 0.00013 | 0.49931 | 7.04968 | 9.50186 | -2.51713 |
| TB 2 | 0.00034 | 0.49804 | -5.02632 | 9.50465 | -2.48571 |
| TB 3 | 0.00023 | 0.49916 | -5.00322 | 9.50323 | 2.46440 |
| TB 4 | 0.00004 | 0.49993 | 7.01744 | 9.50526 | 2.50370 |

ANGLE OF LARGE TRIANGLE = 37.25°

# Text boxes

TB 4

TB 3

TB 2

+Z

+Y

+X

-1.000” OFFSET -X

(Solid Green)

+Z

+X

+Y

+Y

TB 1

+Z

+X