LCLS BC 1 Magnet (Modified Poles)

FIDUCIALIZATION REPORT

| Inspector: | Keith Caban |
| --- | --- |
| Responsible Engineer: | Scott DeBarger |
| Date: | Thursday, October 25, 2007 |
| Work Order/Charge No.: | 12-7100-8 |
| Serial Number: | LCLS - 001076 MAG SN – 42608-3 |
| URL of Fiducial Report: | \\Web002\www-group\met\Quality\FIDUCIAL REPORTS\LCLS BC1 MAGNETS\001076-10/25/07.pdf |

Part Set-up – Coordinate System Set-up

** Poles have been modified

Planar Alignment

- Mid-Plane of the magnet

Spatial Alignment

  - Top of magnet where label is.

“Z” Zero

- Mid-Plane of the magnet.

“X” & “Y” Zero

- X zero is created by the symmetry point of the inner stainless steel where coils lie.

    - Both ends of both sides.

      - Create lines which is y axis.

        - This Creates X- zero

- Symmetry of pole planes

  - This creates Y- zero

Tooling Ball Measurements/Locations

| Tooling Ball | FORM | DIAMETER | X | Y | Z |
| --- | --- | --- | --- | --- | --- |
| TB 1 | 0.00079 | 0.49632 | 6.04474 | 5.65981 | 2.52807 |
| TB 2 | 0.00095 | 0.49412 | 6.04650 | 5.65856 | -2.52780 |
| TB 3 | 0.00036 | 0.49863 | -6.04434 | 5.67337 | -2.53692 |
| TB 4 | 0.00128 | 0.49442 | -6.04895 | 5.67816 | 2.54272 |

| GAP | GAP DIST | X | Z |
| --- | --- | --- | --- |
| 1 | 1.69183 | -2.96370 | 2.33534 |
| 2 | 1.69077 | 0.02298 | -0.10091 |
| 3 | 1.69189 | 2.96343 | 2.33532 |
| 4 | 1.67170 | -2.96354 | -2.33577 |
| 5 | 1.69083 | 2.96353 | -0.33586 |
| 6 | 1.69114 | -0.49816 | -1.41421 |

** Gap location are related to fiducial coordinate system

Nominal Gap per SA-380-321-62 = 1.6920 +0.0005

Distance 7.087 ± 0.0005 measures 7.08715 & 7.08690

# Text boxes

+Z

+X

+Y

+Y

+Z

+X