LCLS BFW WIRECARD ASSY

INSPECTION/FIDUCIAL REPORT

| Inspector: | Keith Caban |
| --- | --- |
| Responsible Engineer: | Venkat Srinvivasan |
| Date: | Tuesday, May 20, 2008 Revised Wednesday, September 24, 2008 |
| Work Order/Charge No.: | 23800-1 |
| Serial Number: | BFW – 04 / VC – 04 / WC – 02 / BASE – 09 |
| Units of Measure: | Metric – Millimeters (mm.) |
| URL | http://www-group/met/Quality/FIDUCIAL%20REPORTS/LCLS%20BFW/BFW-04_VC-04_WC-02_BAS-09.pdf |

Coordinate System Setups

Bench CSY (BENCH$CSY)

(The only CSY in Both States of measurement - out and under vacuum)

- Plane measurement on top of Bench to set zero in Y and is the Spatial Alignment.

- Measure Pin and Diamond Pin to get centers, create a line between the two center points and this sets zero in Z and is the Planar Alignment.

- Take the symmetry point of the two pins this sets zero in X.

Bench Tooling Ball CSY (BTB$CSY)

- Create a plane using Tooling Balls 1, 2, & 3

  - This sets Zero in Y and is the spatial alignment

- Create an axis/line using TB’s 2 to 1 (+X)

  - This is the planar alignment

- TB 2 is a single point that sets zero in X, Y, Z.

BFW Tooling Ball CSY (BFWTB$CSY)

- Create a plane using Tooling Balls 5, 6, & 7

  - This sets Zero in Y and is the spatial alignment

- Create an axis/line using TB’s 6 to 5 (+X)

  - This is the planar alignment

- TB 6 is a single point that sets zero in X, Y, Z.

| View from Z+ or Downstream | View from X+ Side |
| --- | --- |

Wire CSY (WIRE$CSY)

- Back Plane (-Z) Side sets Zero in Z and is the Spatial Alignment

- Looking from +Z Side, Bottom of Card Sets Zero is the Planar Alignment

- The Intersection of the Left Side of the Card and the Bottom, sets Zero in X & Y

Wire offset from Edges per T. Nakashima’s Optical Measurement

| Wire | Wire Location | Wire Diameter |
| --- | --- | --- |
| Vertical Wire (X Location from Card) | 39.6258 | 0.0393 |
| Horizontal Wire (Y Location from Card) | 37.0215 | 0.0407 |

Support Stand CSY (BASE$CSY)

- A plane on Torque Fixture sets Zero in Y and is the Spatial Alignment

- The symmetry of the outer surfaces where “X” fine adjustments are sets zero in X.

- A line where the Torque Fixture 2X Dowel pins sets zero in Z and is the Planar Alignment.

Beamline CSY (BEAMLINE$CSY)

- Use the BENCH$CSY setup except use the Torque Fixture plate to set zero in Y and the Spatial Alignment, transform +258.67 in Y to replicate Theoretical Beamline Center

- This coordinate system will tell us where the Wire Crosshairs are to a theoretical Beam Centerline.

- The wires should be offset ~ -0.400 in X, -0.200 in Y.

Statistical Data (mm.)

Out of Vacuum

100 Cycles (36 recordings) Used WC 00 for out of vacuum testing.

Card Locations  
CSY = BENCH$CSY

| STATS | X-POS CARD | X-INTPT CARD | Y-INTPT CARD |
| --- | --- | --- | --- |
| AVG. | 0.02767 | -39.96180 | -94.86088 |
| Range | 0.19238 | 0.21326 | 0.02873 |
| St. Dev. | 0.05463 | 0.06009 | 0.00780 |

Wire Location

Initially Not Calculated/Measured on this BFW, Measurements/Stats are done from MMF CMM Cycle of 10 by Ed Reese (09/24/08, No Vacuum).

CSY = BEAMLINE$CSY

| Stats | X | Y |
| --- | --- | --- |
| AVG. | -0.54062 | -0.17020 |
| Range | 0.11740 | 0.00610 |
| St. Dev. | 0.03410 | 0.00171 |
| Nom. | -0.40000 | -0.20000 |
| RMS Tol | 0.06500 | 0.03000 |

Transformation WIRE$CSY TO BENCH$CSY

Not Calculated/Measured on this BFW (Not part of inspection at this point)

| Stats | X | Y | Z | APA_PXPY | PA_PYPZ | PA_PXPZ |
| --- | --- | --- | --- | --- | --- | --- |
| AVG. |  |  |  |  |  |  |
| Range |  |  |  |  |  |  |
| St. Dev. |  |  |  |  |  |  |
| RMS Tol | 0.06500 | 0.03000 | - | - | - | - |

Transformation WIRE$CSY to BFWTB$CSY

Not Calculated/Measured on this BFW (Not part of inspection at this point)

| Stats | X | Y | Z | APA_PXPY | PA_PYPZ | PA_PXPZ |
| --- | --- | --- | --- | --- | --- | --- |
| AVG. |  |  |  |  |  |  |
| Range |  |  |  |  |  |  |
| St. Dev. |  |  |  |  |  |  |
| RMS Tol | 0.06500 | 0.03000 | - | - | - | - |

Transformation WIRE$CSY to BFWTB$CSY (The Actual Transformation used on Both CMM’s

| X | Y | Z | APA_PXPY | PA_PYPZ | PA_PXPZ |
| --- | --- | --- | --- | --- | --- |
| 94.01900 | -160.35774 | 35.78696 | - | - | - |

Under Vacuum

10 Cycles (10 recordings)

TB 5-8 Locations

CSY = BENCH$CSY

| TB 5 | X | Y | Z |
| --- | --- | --- | --- |
| AVG. | 94.26061 | 102.58573 | -38.22970 |
| Range | 0.01059 | 0.01408 | 0.00573 |
| St. Dev. | 0.00349 | 0.00438 | 0.00170 |
| RMS Tol | 0.06500 | 0.03000 | - |

| TB 6 | X | Y | Z |
| --- | --- | --- | --- |
| AVG. | -94.60578 | 102.44049 | -37.96454 |
| Range | 0.01131 | 0.01855 | 0.00628 |
| St. Dev. | 0.00368 | 0.00598 | 0.00217 |
| RMS Tol | 0.06500 | 0.03000 | - |

| TB 7 | X | Y | Z |
| --- | --- | --- | --- |
| AVG. | -51.33575 | 102.47507 | 67.80879 |
| Range | 0.00884 | 0.00498 | 0.00540 |
| St. Dev. | 0.00282 | 0.00152 | 0.00183 |
| RMS Tol | 0.06500 | 0.03000 | - |

| TB 8 | X | Y | Z |
| --- | --- | --- | --- |
| AVG. | 51.26168 | 102.57015 | 67.67853 |
| Range | 0.00890 | 0.01270 | 0.00487 |
| St. Dev. | 0.00286 | 0.00445 | 0.00160 |
| RMS Tol | 0.06500 | 0.03000 | - |

Wire Location

CSY = BEAMLINE$CSY

Not Calculated/Measured on this BFW (Not part of inspection at this point)

| Stats | X | Y |
| --- | --- | --- |
| AVG. |  |  |
| Range |  |  |
| St. Dev. |  |  |
| Nom. | -0.40000 | -0.20000 |
| RMS Tol | 0.06500 | 0.03000 |

Transformation WIRE$CSY TO BENCH$CSY

Not Calculated/Measured on this BFW (Not part of inspection at this point)

| Stats | X | Y | Z | APA_PXPY | PA_PYPZ | PA_PXPZ |
| --- | --- | --- | --- | --- | --- | --- |
| AVG. |  |  |  |  |  |  |
| Range |  |  |  |  |  |  |
| St. Dev. |  |  |  |  |  |  |
| RMS Tol | 0.06500 | 0.03000 | - | - | - | - |

Locations of ¼-20 Jo-Plugs Cylinders Intersected into Plane projected 12.7mm away from surface replicating a ½” (12.7mm) offset to center of Tooling Ball.

CSY = BEAMLINE$CSY

Not Calculated/Measured on this BFW (Not part of inspection at this point)

| INTPT | X | Y | Z |
| --- | --- | --- | --- |
| 1 |  |  |  |
| 2 |  |  |  |
| 3 |  |  |  |
| 4 |  |  |  |

# Text boxes

Y+

+Z

Z+

Y+

+X

+Y

X+

Transformation