LCLS BFW WIRECARD ASSY

FIDUCIAL REPORT

| Inspector: | Keith Caban |
| --- | --- |
| Responsible Engineer: | Venkat Srinivasan |
| Date: | Monday, May 12, 2008 |
| Work Order/Charge No.: | 23800-1 |
| Serial Number: | BFW – 04 / WC – 02 / BASE – 09 |
| URL | http://www-group/met/Quality/FIDUCIAL%20REPORTS/LCLS%20BFW/BFW-04_WC-02.pdf |

Basic Procedure:

- Locate Bench and Bench CSY

- Install Top Portion of BFW

- Mount Wire Card and Center to Bench CSY

- 5 Cycles of Repeatability to Bench CSY of

  - Positions of Wire Card (Center, 2 Corners/Intersection Points)

  - 4 – BFW Fiducials (X,Y,Z locations)

- Create Wire Card CSY

- Transform to Wire Cross Hairs (Optical CMM Location Offset to Outer Datums of Card)

- Translate 4 – BFW Fiducials into Wire CSY

- Use 4X Screws to Secure the BFW in its Up or “In-Beam Position” for Transport and Girder Alignment

- Place BFW Chamber onto Support Stand, Set to Nominal Dimensions in X (0.000”) & Y (10.184”).

- Place Top portion of BFW onto BFW Chamber and Support Stand

  - Tighten down per ANL Torque Sequence/Specs

  - Coordinate with MFD Vacuum for Pump Down

Coordinate System Setups

Bench CSY

- Plane measurement on top of Bench to set zero in Y and is the Spatial Alignment.

- Measure Pin and Diamond Pin to get centers, create a line between the two center points and this sets zero in Z and is the Planar Alignment.

- Take the symmetry point of the two pins this sets zero in X.

Wire CSY

| View from Z+ or Downstream | View from X+ Side |
| --- | --- |

- Back Plane (-Z) Side sets Zero in Z and is the Spatial Alignment

- Looking from +Z Side, Bottom of Card Sets Zero is the Planar Alignment

- The Intersection of the Left Side of the Card and the Bottom, sets Zero in X & Y

Shift this CSY using the Optical CMM Output for the exact offsets (file in Sharepoint SV-VS122108-Wire.xls per Serialized Wire Card.

Transform TB’s into Wire CSY for Fiducial Numbers

**See Keith Caban or Venkat Srinivasan for Repeatability Data

Initial Repeatability Data 100 cycles used WC 00, switched after putting tabs on to Wirecard to WC-02

RMS Data

Out of Vacuum

| Wirecard | SYM X | INT X | INT Y |
| --- | --- | --- | --- |
| 00 | 0.05392 | 0.05929 | 0.00772 |

| TB | X | Y | Z |
| --- | --- | --- | --- |
| 5 | 0.00904 | 0.02075 | 0.00644 |
| 6 | 0.00892 | 0.02527 | 0.00616 |
| 7 | 0.01573 | 0.00285 | 0.00317 |
| 8 | 0.01572 | 0.02417 | 0.00364 |

Under Vacuum

| TB | X | Y | Z |
| --- | --- | --- | --- |
| 5 | 0.00777 | 0.01085 | 0.00439 |
| 6 | 0.00802 | 0.01239 | 0.00428 |
| 7 | 0.00656 | 0.00397 | 0.00407 |
| 8 | 0.00668 | 0.00911 | 0.00367 |

Fiducial Data

CSY : Wire$CSY (Wire Crosshair)

Wire Card is positioned in X to -0.01848 microns

Wire Card Offsets for WC - 02

(39.600) Vertical =		39.6258

(36.997) Horizontal =	37.0215

| TB | Form | Diameter | X | Y | Z |
| --- | --- | --- | --- | --- | --- |
| 5 | 0.00471 | 12.68243 | 94.94689 | 160.44747 | -35.07872 |
| 6 | 0.00117 | 12.67051 | -93.90597 | 160.51473 | -35.37761 |
| 7 | 0.00159 | 12.68254 | -50.96382 | 160.13127 | 70.51473 |
| 8 | 0.00723 | 12.64781 | 51.62755 | 160.12347 | 70.69432 |

BFW Support Assembly Alignment

Support Stand Tolerance

X ± 0.060mm or ± 0.0024

Y ± 0.040mm or ± 0.0016

Support Stand SN = 09

Chamber (Waveguide) Thickness = 0.263

Distance to the Top Flange (mating) = 12.46619

Position Y (10.1841) (Beamline) = 10.1840

Position X (0.0000) = -0.00130

Bottom of Form

# Text boxes

Z+

+Z

X+

Y+

+X

+Y

Transformation

Y+