LCLS BFW WIRECARD ASSY

INSPECTION/FIDUCIAL REPORT

| Inspector: | Keith Caban |
| --- | --- |
| Responsible Engineer: | Venkat Srinvivasan |
| Date: | Tuesday, May 20, 2008 Revised Tuesday, September 23, 2008 |
| Work Order/Charge No.: | 23800-1 |
| Serial Number: | BFW – 05 / VC – 05 / WC – 05 / BASE – 14 |
| Units of Measure: | Metric – Millimeters (mm.) |
| URL | http://www-group/met/Quality/FIDUCIAL%20REPORTS/LCLS%20BFW/BFW-05_VC-05_WC-05_BAS-14.pdf |

Coordinate System Setups

Bench CSY (BENCH$CSY)

(The only CSY in Both States of measurement - out and under vacuum)

- Plane measurement on top of Bench to set zero in Y and is the Spatial Alignment.

- Measure Pin and Diamond Pin to get centers, create a line between the two center points and this sets zero in Z and is the Planar Alignment.

- Take the symmetry point of the two pins this sets zero in X.

Bench Tooling Ball CSY (BTB$CSY)

- Create a plane using Tooling Balls 1, 2, & 3

  - This sets Zero in Y and is the spatial alignment

- Create an axis/line using TB’s 2 to 1 (+X)

  - This is the planar alignment

- TB 2 is a single point that sets zero in X, Y, Z.

BFW Tooling Ball CSY (BFWTB$CSY)

- Create a plane using Tooling Balls 5, 6, & 7

  - This sets Zero in Y and is the spatial alignment

- Create an axis/line using TB’s 6 to 5 (+X)

  - This is the planar alignment

- TB 6 is a single point that sets zero in X, Y, Z.

| View from Z+ or Downstream | View from X+ Side |
| --- | --- |

Wire CSY (WIRE$CSY)

- Back Plane (-Z) Side sets Zero in Z and is the Spatial Alignment

- Looking from +Z Side, Bottom of Card Sets Zero is the Planar Alignment

- The Intersection of the Left Side of the Card and the Bottom, sets Zero in X & Y

Wire offset from Edges per T. Nakashima’s Optical Measurement

| Wire | Wire Location | Wire Diameter |
| --- | --- | --- |
| Vertical Wire (X Location from Card) | 39.6203 | 0.0396 |
| Horizontal Wire (Y Location from Card) | 37.0005 | 0.0388 |

Support Stand CSY (BASE$CSY)

- A plane on Torque Fixture sets Zero in Y and is the Spatial Alignment

- The symmetry of the outer surfaces where “X” fine adjustments are sets zero in X.

- A line where the Torque Fixture 2X Dowel pins sets zero in Z and is the Planar Alignment.

Beamline CSY (BEAMLINE$CSY)

- Use the BENCH$CSY setup except use the Torque Fixture plate to set zero in Y and the Spatial Alignment, transform +258.67 in Y to replicate Theoretical Beamline Center

- This coordinate system will tell us where the Wire Crosshairs are to a theoretical Beam Centerline.

- The wires should be offset ~ -0.400 in X, -0.200 in Y.

Statistical Data (mm.)

Out of Vacuum

10 Cycles (10 recordings)

Card Locations  
CSY = BENCH$CSY

| STATS | X-POS CARD | X-INTPT CARD | Y-INTPT CARD |
| --- | --- | --- | --- |
| AVG. | -0.17960 |  | -95.26975 |
| Range | 0.18770 |  | 0.02724 |
| St. Dev. | 0.05453 |  | 0.00751 |

Wire Location

Initially Not Calculated/Measured on this BFW, Measurements/Stats are done from MMF CMM Cycle of 10 by Ed Reese (09/23/08, No Vacuum).

CSY = BEAMLINE$CSY

| Stats | X | Y |
| --- | --- | --- |
| AVG. | -0.48704 | -0.16272 |
| Range | 0.07320 | 0.00260 |
| St. Dev. | 0.02167 | 0.00086 |
| Nom. | -0.40000 | -0.20000 |
| RMS Tol | 0.06500 | 0.03000 |

Transformation WIRE$CSY TO BENCH$CSY

Not Calculated/Measured on this BFW (Not part of inspection at this point)

| Stats | X | Y | Z | APA_PXPY | PA_PYPZ | PA_PXPZ |
| --- | --- | --- | --- | --- | --- | --- |
| AVG. |  |  |  |  |  |  |
| Range |  |  |  |  |  |  |
| St. Dev. |  |  |  |  |  |  |
| RMS Tol | 0.06500 | 0.03000 | - | - | - | - |

Transformation WIRE$CSY to BFWTB$CSY

| Stats | X | Y | Z | APA_PXPY | PA_PYPZ | PA_PXPZ |
| --- | --- | --- | --- | --- | --- | --- |
| AVG. | 94.06780 | -160.55627 | 35.84900 | -0.06129 | 89.80929 | 90.08003 |
| Range | 0.00501 | 0.00259 | 0.00680 | 0.00960 | 0.00432 | 0.00614 |
| St. Dev. | 0.00148 | 0.00078 | 0.00194 | 0.00335 | 0.00131 | 0.00202 |
| RMS Tol | 0.06500 | 0.03000 | - | - | - | - |

Under Vacuum

100 Cycles (19 recordings)

TB 5-8 Locations

CSY = BENCH$CSY

| TB 5 | X | Y | Z |
| --- | --- | --- | --- |
| AVG. | 94.69571 | 101.77761 | -37.84015 |
| Range | 0.02387 | 0.05760 | 0.07799 |
| St. Dev. | 0.00990 | 0.02328 | 0.01609 |
| RMS Tol | 0.06500 | 0.03000 | - |

| TB 6 | X | Y | Z |
| --- | --- | --- | --- |
| AVG. | -94.18932 | 102.26439 | -38.11829 |
| Range | 0.02376 | 0.06269 | 0.03070 |
| St. Dev. | 0.01003 | 0.02588 | 0.00829 |
| RMS Tol | 0.06500 | 0.03000 | - |

| TB 7 | X | Y | Z |
| --- | --- | --- | --- |
| AVG. | -51.21688 | 102.41837 | 67.78256 |
| Range | 0.05890 | 0.03087 | 0.00834 |
| St. Dev. | 0.01889 | 0.00669 | 0.00319 |
| RMS Tol | 0.06500 | 0.03000 | - |

| TB 8 | X | Y | Z |
| --- | --- | --- | --- |
| AVG. | 51.36915 | 102.16770 | 67.93734 |
| Range | 0.05914 | 0.05571 | 0.05202 |
| St. Dev. | 0.01900 | 0.02362 | 0.01072 |
| RMS Tol | 0.06500 | 0.03000 | - |

Wire Location

CSY = BEAMLINE$CSY

Not Calculated/Measured on this BFW (Not part of inspection at this point)

| Stats | X | Y |
| --- | --- | --- |
| AVG. |  |  |
| Range |  |  |
| St. Dev. |  |  |
| Nom. | -0.40000 | -0.20000 |
| RMS Tol | 0.06500 | 0.03000 |

Transformation WIRE$CSY TO BENCH$CSY

| Stats | X | Y | Z | APA_PXPY | PA_PYPZ | PA_PXPZ |
| --- | --- | --- | --- | --- | --- | --- |
| AVG. | -0.59190 | -58.45328 | -1.70882 | -0.21704 | 89.66425 | 90.18825 |
| Range | 0.12873 | 0.00632 | 0.04771 | 0.03572 | 0.02332 | 0.03302 |
| St. Dev. | 0.05344 | 0.00251 | 0.01851 | 0.01488 | 0.00731 | 0.00723 |
| RMS Tol | 0.06500 | 0.03000 | - | - | - | - |

Locations of ¼-20 Jo-Plugs Cylinders Intersected into Plane projected 12.7mm away from surface replicating a ½” (12.7mm) offset to center of Tooling Ball.

CSY = BEAMLINE$CSY

Not Calculated/Measured on this BFW (Not part of inspection at this point)

| INTPT | X | Y | Z |
| --- | --- | --- | --- |
| 1 |  |  |  |
| 2 |  |  |  |
| 3 |  |  |  |
| 4 |  |  |  |

# Text boxes

Y+

+Z

Z+

Y+

+X

+Y

X+

Transformation