LCLS BFW WIRECARD ASSY

INSPECTION/FIDUCIAL REPORT

| Inspector: | Keith Caban |
| --- | --- |
| Responsible Engineer: | Venkat Srinvivasan |
| Date: | Thursday, May 29, 2008 Revised Thursday, September 25, 2008 |
| Work Order/Charge No.: | 23800-1 |
| Serial Number: | BFW – 06 / VC – 06 / WC – 38 / BASE – 19 |
| Units of Measure: | Metric – Millimeters (mm.) |
| URL | http://www-group/met/Quality/FIDUCIAL%20REPORTS/LCLS%20BFW/BFW-06_VC-06_WC-38_BAS-19.pdf |

Coordinate System Setups

Bench CSY (BENCH$CSY)

(The only CSY in Both States of measurement - out and under vacuum)

- Plane measurement on top of Bench to set zero in Y and is the Spatial Alignment.

- Measure Pin and Diamond Pin to get centers, create a line between the two center points and this sets zero in Z and is the Planar Alignment.

- Take the symmetry point of the two pins this sets zero in X.

Bench Tooling Ball CSY (BTB$CSY)

- Create a plane using Tooling Balls 1, 2, & 3

  - This sets Zero in Y and is the spatial alignment

- Create an axis/line using TB’s 2 to 1 (+X)

  - This is the planar alignment

- TB 2 is a single point that sets zero in X, Y, Z.

BFW Tooling Ball CSY (BFWTB$CSY)

- Create a plane using Tooling Balls 5, 6, & 7

  - This sets Zero in Y and is the spatial alignment

- Create an axis/line using TB’s 6 to 5 (+X)

  - This is the planar alignment

- TB 6 is a single point that sets zero in X, Y, Z.

| View from Z+ or Downstream | View from X+ Side |
| --- | --- |

Wire CSY (WIRE$CSY)

- Back Plane (-Z) Side sets Zero in Z and is the Spatial Alignment

- Looking from +Z Side, Bottom of Card Sets Zero is the Planar Alignment

- The Intersection of the Left Side of the Card and the Bottom, sets Zero in X & Y

Wire offset from Edges per T. Nakashima’s Optical Measurement

| Wire | Wire Location | Wire Diameter |
| --- | --- | --- |
| Vertical Wire (X Location from Card) | 39.6181 | 0.0366 |
| Horizontal Wire (Y Location from Card) | 37.0054 | 0.0365 |

Support Stand CSY (BASE$CSY)

- A plane on Torque Fixture sets Zero in Y and is the Spatial Alignment

- The symmetry of the outer surfaces where “X” fine adjustments are sets zero in X.

- A line where the Torque Fixture 2X Dowel pins sets zero in Z and is the Planar Alignment.

Beamline CSY (BEAMLINE$CSY)

- Use the BENCH$CSY setup except use the Torque Fixture plate to set zero in Y and the Spatial Alignment, transform +258.67 in Y to replicate Theoretical Beamline Center

- This coordinate system will tell us where the Wire Crosshairs are to a theoretical Beam Centerline.

- The wires should be offset ~ -0.400 in X, -0.200 in Y.

Statistical Data (mm.)

Out of Vacuum

100 Cycles (20 recordings)

Card Locations  
CSY = BENCH$CSY

| STATS | X-POS CARD | X-INTPT CARD | Y-INTPT CARD |
| --- | --- | --- | --- |
| AVG. | -0.09725 | -40.29089 | -96.93552 |
| Range | 0.05948 | 0.09160 | 0.22572 |
| St. Dev. | 0.02090 | 0.03136 | 0.08479 |

Wire Location

Initially Not Calculated/Measured on this BFW, Measurements/Stats are done from MMF CMM Cycle of 10 by Ed Reese (09/25/08, No Vacuum).

CSY = BEAMLINE$CSY

| Stats | X | Y |
| --- | --- | --- |
| AVG. | -0.57021 | -0.22564 |
| Range | 0.00520 | 0.01150 |
| St. Dev. | 0.00166 | 0.00347 |
| Nom. | -0.40000 | -0.20000 |
| RMS Tol | 0.06500 | 0.03000 |

Transformation WIRE$CSY TO BENCH$CSY

| Stats | X | Y | Z | APA_PXPY | PA_PYPZ | PA_PXPZ |
| --- | --- | --- | --- | --- | --- | --- |
| AVG. | -0.51399 | -60.12884 | -2.86736 | -0.28130 | 89.98078 | 89.92737 |
| Range | 0.06073 | 0.22526 | 0.11724 | 0.01021 | 0.03742 | 0.06978 |
| St. Dev. | 0.02133 | 0.08458 | 0.03778 | 0.00294 | 0.01311 | 0.02606 |
| RMS Tol | 0.06500 | 0.03000 | - | - | - | - |

Transformation WIRE$CSY to BFWTB$CSY

| Stats | X | Y | Z | APA_PXPY | PA_PYPZ | PA_PXPZ |
| --- | --- | --- | --- | --- | --- | --- |
| AVG. | 94.61569 | -160.39996 | 35.64864 | -0.07569 | 89.86374 | 89.91773 |
| Range | 0.00485 | 0.01781 | 0.02760 | 0.00874 | 0.01422 | 0.01376 |
| St. Dev. | 0.00120 | 0.00476 | 0.00686 | 0.00214 | 0.00383 | 0.00346 |
| RMS Tol | 0.06500 | 0.03000 | - | - | - | - |

Under Vacuum

200 Cycles (22 recordings)

TB 5-8 Locations

CSY = BENCH$CSY

| TB 5 | X | Y | Z |
| --- | --- | --- | --- |
| AVG. | 94.35345 | 99.89047 | -38.00066 |
| Range | 0.02112 | 0.05318 | 0.02755 |
| St. Dev. | 0.00692 | 0.01749 | 0.00827 |
| RMS Tol | 0.06500 | 0.03000 | - |

| TB 6 | X | Y | Z |
| --- | --- | --- | --- |
| AVG. | -94.51201 | 100.72217 | -38.16673 |
| Range | 0.02155 | 0.06686 | 0.08398 |
| St. Dev. | 0.00722 | 0.02302 | 0.02939 |
| RMS Tol | 0.06500 | 0.03000 | - |

| TB 7 | X | Y | Z |
| --- | --- | --- | --- |
| AVG. | -51.47475 | 100.51288 | 67.71628 |
| Range | 0.01560 | 0.09637 | 0.06926 |
| St. Dev. | 0.00530 | 0.03353 | 0.02412 |
| RMS Tol | 0.06500 | 0.03000 | - |

| TB 8 | X | Y | Z |
| --- | --- | --- | --- |
| AVG. | 51.09411 | 100.07531 | 67.79779 |
| Range | 0.01584 | 0.08831 | 0.03894 |
| St. Dev. | 0.00540 | 0.03028 | 0.01278 |
| RMS Tol | 0.06500 | 0.03000 | - |

Wire Location

CSY = BEAMLINE$CSY

Not Calculated/Measured on this BFW (Not part of inspection at this point)

| Stats | X | Y |
| --- | --- | --- |
| AVG. |  |  |
| Range |  |  |
| St. Dev. |  |  |
| Nom. | -0.40000 | -0.20000 |
| RMS Tol | 0.06500 | 0.03000 |

Transformation WIRE$CSY TO BENCH$CSY

| Stats | X | Y | Z | APA_PXPY | PA_PYPZ | PA_PXPZ |
| --- | --- | --- | --- | --- | --- | --- |
| AVG. | -0.63461 | -60.09341 | -2.45651 | -0.32926 | 89.87056 | 89.97055 |
| Range | 0.00723 | 0.07090 | 0.00429 | 0.00559 | 0.01846 | 0.02030 |
| St. Dev. | 0.00213 | 0.02421 | 0.00106 | 0.00180 | 0.00640 | 0.00686 |
| RMS Tol | 0.06500 | 0.03000 | - | - | - | - |

Locations of ¼-20 Jo-Plugs Cylinders Intersected into Plane projected 12.7mm away from surface replicating a ½” (12.7mm) offset to center of Tooling Ball.

CSY = BEAMLINE$CSY

Not Calculated/Measured on this BFW (Not part of inspection at this point)

| INTPT | X | Y | Z |
| --- | --- | --- | --- |
| 1 |  |  |  |
| 2 |  |  |  |
| 3 |  |  |  |
| 4 |  |  |  |

# Text boxes

Y+

+Z

Z+

Y+

+X

+Y

X+

Transformation