LCLS BFW WIRECARD ASSY

FIDUCIAL REPORT

| Inspector: | Keith Caban |
| --- | --- |
| Responsible Engineer: | Venkat Srinivasan |
| Date: | Thursday, May 29, 2008 |
| Work Order/Charge No.: | 23800-1 |
| Serial Number: | BFW – 06 / WC – 38 / BASE – 19 |
| URL | http://www-group/met/Quality/FIDUCIAL%20REPORTS/LCLS%20BFW/BFW-06_WC-38_BAS-19.pdf |

Coordinate System Setups

Bench CSY (BENCH$CSY)

- Plane measurement on top of Bench to set zero in Y and is the Spatial Alignment.

- Measure Pin and Diamond Pin to get centers, create a line between the two center points and this sets zero in Z and is the Planar Alignment.

- Take the symmetry point of the two pins this sets zero in X.

Bench Tooling Ball CSY (BTB$CSY)

- Create a plane using Tooling Balls 1, 2, & 3

  - This sets Zero in Y and is the spatial alignment

- Create an axis/line using TB’s 2 to 1 (+X)

  - This is the planar alignment

- TB 2 is a single point that sets zero in X, Y, Z.

BFW Tooling Ball CSY (BFWTB$CSY)

- Create a plane using Tooling Balls 5, 6, & 7

  - This sets Zero in Y and is the spatial alignment

- Create an axis/line using TB’s 6 to 5 (+X)

  - This is the planar alignment

- TB 6 is a single point that sets zero in X, Y, Z.

| View from Z+ or Downstream | View from X+ Side |
| --- | --- |

Wire CSY (WIRE$CSY)

- Back Plane (-Z) Side sets Zero in Z and is the Spatial Alignment

- Looking from +Z Side, Bottom of Card Sets Zero is the Planar Alignment

- The Intersection of the Left Side of the Card and the Bottom, sets Zero in X & Y

Wire offset from Edges per T. Nakashima’s Optical Measurement

| Vertical Wire (X Location from Card) | 39.6181 |
| --- | --- |
| Horizontal Wire (Y Location from Card) | 37.0054 |

Support Stand CSY (BASE$CSY)

- A plane on Torque Fixture sets Zero in Y and is the Spatial Alignment

- The symmetry of the outer surfaces where “X” fine adjustments are sets zero in X.

- A line where the Torque Fixture 2X Dowel pins sets zero in Z and is the Planar Alignment.

Statistical Data

Out of Vacuum

Card Locations  
CSY = BENCH$CSY

| STATS | X-POS CARD | X-INTPT CARD | Y-INTPT CARD |
| --- | --- | --- | --- |
| average | -0.09725 | -40.29089 | -96.93552 |
| range | 0.05948 | 0.09160 | 0.22572 |
| st dev | 0.02090 | 0.03136 | 0.08479 |

Out of Vacuum Cont…

TB 5-8 Locations

CSY = BTB$CSY

| TB 5 | X | Y | Z |
| --- | --- | --- | --- |
| AVG. | 226.84307 | 87.24812 | -95.70151 |
| Range | 0.10111 | 0.21170 | 0.06630 |
| St. Dev. | 0.03606 | 0.07816 | 0.02609 |

| TB 6 | X | Y | Z |
| --- | --- | --- | --- |
| AVG. | 37.97104 | 87.94544 | -95.73015 |
| Range | 0.10284 | 0.21390 | 0.30167 |
| St. Dev. | 0.03692 | 0.08037 | 0.11113 |

| TB 7 | X | Y | Z |
| --- | --- | --- | --- |
| AVG. | 81.08936 | 87.58584 | 10.11687 |
| Range | 0.03711 | 0.27521 | 0.24925 |
| St. Dev. | 0.01195 | 0.10255 | 0.09199 |

| TB 8 | X | Y | Z |
| --- | --- | --- | --- |
| AVG. | 183.66119 | 87.21981 | 10.12226 |
| Range | 0.03724 | 0.25614 | 0.12425 |
| St. Dev. | 0.01200 | 0.09761 | 0.04698 |

Transformation BFWTB$CSY to BTB$CSY

| Stats | X | Y | Z | APA_PXPY | PA_PYPZ | PA_PXPZ |
| --- | --- | --- | --- | --- | --- | --- |
| AVG. | 37.97104 | 87.94544 | -95.73015 | -0.21154 | 90.10846 | 90.00908 |
| Range | 0.10284 | 0.21390 | 0.30167 | 0.01253 | 0.04083 | 0.07144 |
| St. Dev. | 0.03692 | 0.08037 | 0.11113 | 0.00400 | 0.01334 | 0.02622 |

Transformation WIRE$CSY to BTB$CSY

| Stats | X | Y | Z | APA_PXPY | PA_PYPZ | PA_PXPZ |
| --- | --- | --- | --- | --- | --- | --- |
| AVG. | 131.98826 | -72.86993 | -60.37095 | -0.28707 | 89.96739 | 89.92285 |
| Range | 0.06074 | 0.22524 | 0.11728 | 0.01021 | 0.11980 | 0.06977 |
| St. Dev. | 0.02134 | 0.08457 | 0.03779 | 0.00294 | 0.02645 | 0.02570 |

Transformation WIRE$CSY to BFWTB$CSY

| Stats | X | Y | Z | APA_PXPY | PA_PYPZ | PA_PXPZ |
| --- | --- | --- | --- | --- | --- | --- |
| AVG. | 94.71453 | -160.61536 | 35.89852 | 0.08987 | 89.75101 | 90.03282 |
| Range | 0.00687 | 0.00347 | 0.05115 | 0.00304 | 0.00903 | 0.04439 |
| St. Dev. | 0.00148 | 0.00098 | 0.01033 | 0.00069 | 0.00266 | 0.00937 |

Transformation WIRE$CSY to BENCH$CSY

| Stats | X | Y | Z | APA_PXPY | PA_PYPZ | PA_PXPZ |
| --- | --- | --- | --- | --- | --- | --- |
| AVG. | -0.51399 | -60.12884 | -2.86736 | -0.28130 | 89.98078 | 89.92737 |
| Range | 0.06073 | 0.22526 | 0.11724 | 0.01021 | 0.03742 | 0.06978 |
| St. Dev. | 0.02133 | 0.08458 | 0.03778 | 0.00294 | 0.01311 | 0.02606 |

Under Vacuum

TB 5-8 Locations

CSY = BENCH$CSY

| TB 5 | X | Y | Z |
| --- | --- | --- | --- |
| AVG. | 94.35345 | 99.89047 | -38.00066 |
| Range | 0.02112 | 0.05318 | 0.02755 |
| St. Dev. | 0.00692 | 0.01749 | 0.00827 |

| TB 6 | X | Y | Z |
| --- | --- | --- | --- |
| AVG. | -94.51201 | 100.72217 | -38.16673 |
| Range | 0.02155 | 0.06686 | 0.08398 |
| St. Dev. | 0.00722 | 0.02302 | 0.02939 |

| TB 7 | X | Y | Z |
| --- | --- | --- | --- |
| AVG. | -51.47475 | 100.51288 | 67.71628 |
| Range | 0.01560 | 0.09637 | 0.06926 |
| St. Dev. | 0.00530 | 0.03353 | 0.02412 |

| TB 8 | X | Y | Z |
| --- | --- | --- | --- |
| AVG. | 51.09411 | 100.07531 | 67.79779 |
| Range | 0.01584 | 0.08831 | 0.03894 |
| St. Dev. | 0.00540 | 0.03028 | 0.01278 |

Transformation BFWTB$CSY to BASE$CSY

| Stats | X | Y | Z | APA_PXPY | PA_PYPZ | PA_PXPZ |
| --- | --- | --- | --- | --- | --- | --- |
| AVG. | -95.16082 | 417.35635 | -102.14922 | -0.25153 | 90.01283 | 89.96434 |
| Range | 0.02143 | 0.06687 | 0.08401 | 0.00558 | 0.01847 | 0.02032 |
| St. Dev. | 0.00717 | 0.02302 | 0.02940 | 0.00180 | 0.00640 | 0.00686 |

Transformation WIRE$CSY to BENCH$CSY

| Stats | X | Y | Z | APA_PXPY | PA_PYPZ | PA_PXPZ |
| --- | --- | --- | --- | --- | --- | --- |
| AVG. | -0.63461 | -60.09341 | -2.45651 | -0.32926 | 89.87056 | 89.97055 |
| Range | 0.00723 | 0.07090 | 0.00429 | 0.00559 | 0.01846 | 0.02030 |
| St. Dev. | 0.00213 | 0.02421 | 0.00106 | 0.00180 | 0.00640 | 0.00686 |

Locations of ¼-20 Jo-Plugs Cylinders Intersected into Plane

CSY = BFWTB$CSY

| INTPT | X | Y | Z |
| --- | --- | --- | --- |
| A | 67.19944 | 69.15681 | 94.07530 |
| B | 121.22416 | 69.35330 | 94.01670 |
| C | 174.18775 | 22.36993 | -17.72124 |
| D | 14.43508 | 21.67561 | -17.54620 |

Miscellaneous QC Checks

| Checklist | Y/N | Notes |
| --- | --- | --- |
| Out of Crate Damage | N |  |
| Prior to installation, ensure Wire Crosshairs are pertinent | Y |  |
| Wire card SN matches BFW SN | N | WC-06 being used on another BFW, per Venkat, used WC-34 found no vertical wire, so now using WC-38 |
| Card has copper flex wires installed | Y |  |
| Connect Tabs from welded rod to Wire card | Y |  |
| Feedthru Welded Rods nuts tightly secured from 4 Ag nuts per weld rod | Y |  |
| Continuity from Top Feedthru to Card Connection Tabs | Y |  |
| Check 3 places for air fittings seal (RTV or Silicone/Teflon Tape) | Y |  |
| Connect Air @ 90psi, check for air leaks prior to actuation | Y |  |
| During actuation, ensure Limit Switches meet in Target In/Out Positions | Y |  |
| Ensure Light illuminates (LED) when turning on power supply | Y |  |
| Chamber: Look for aluminum coating on inside of chamber | Y |  |
| Chamber: Ensure flange clamp is made of Dichronite (Dark Gray) chain link/standard clamp | Y |  |
| Chamber: Ensure no loose hardware inside vacuum chamber | Y |  |
| Wire Crosshairs check prior to installation for vacuum | Y |  |
| Vacuum Leak Test: NIL (no indicated leak)? | Y |  |
| Ready for Transportation Check: WC is target is “IN” position with 4 holding screws at the top. | Y |  |
| Post Assy Check: Ensure plastic tubing is taped off at the top with mylar/electrical tape | Y |  |

Bottom of Form

# Text boxes

Y+

+Z

Z+

Y+

+X

+Y

X+

Transformation