LCLS BFW WIRECARD ASSY

INSPECTION/FIDUCIAL REPORT

| Inspector: | Keith Caban |
| --- | --- |
| Responsible Engineer: | Venkat Srinivasan |
| Date: | Friday, June 06, 2008 |
| Work Order/Charge No.: | 23800-1 |
| Serial Number: | BFW – 07 / VC – 07 WC – 07 / BASE – 31 |
| Units of Measure: | Metric – Millimeters (mm.) |
| URL | http://www-group/met/Quality/FIDUCIAL%20REPORTS/LCLS%20BFW/BFW-07_VC-07_WC-07_BAS-31.pdf |

Coordinate System Setups

Bench CSY (BENCH$CSY)

(The only CSY in Both States of measurement - out and under vacuum)

- Plane measurement on top of Bench to set zero in Y and is the Spatial Alignment.

- Measure Pin and Diamond Pin to get centers, create a line between the two center points and this sets zero in Z and is the Planar Alignment.

- Take the symmetry point of the two pins this sets zero in X.

Bench Tooling Ball CSY (BTB$CSY)

- Create a plane using Tooling Balls 1, 2, & 3

  - This sets Zero in Y and is the spatial alignment

- Create an axis/line using TB’s 2 to 1 (+X)

  - This is the planar alignment

- TB 2 is a single point that sets zero in X, Y, Z.

BFW Tooling Ball CSY (BFWTB$CSY)

- Create a plane using Tooling Balls 5, 6, & 7

  - This sets Zero in Y and is the spatial alignment

- Create an axis/line using TB’s 6 to 5 (+X)

  - This is the planar alignment

- TB 6 is a single point that sets zero in X, Y, Z.

| View from Z+ or Downstream | View from X+ Side |
| --- | --- |

Wire CSY (WIRE$CSY)

- Back Plane (-Z) Side sets Zero in Z and is the Spatial Alignment

- Looking from +Z Side, Bottom of Card Sets Zero is the Planar Alignment

- The Intersection of the Left Side of the Card and the Bottom, sets Zero in X & Y

Wire offset from Edges per T. Nakashima’s Optical Measurement

| Wire | Wire Location | Wire Diameter |
| --- | --- | --- |
| Vertical Wire (X Location from Card) | 39.6064 | 0.0389 |
| Horizontal Wire (Y Location from Card) | 37.0077 | 0.0397 |

Support Stand CSY (BASE$CSY)

- A plane on Torque Fixture sets Zero in Y and is the Spatial Alignment

- The symmetry of the outer surfaces where “X” fine adjustments are sets zero in X.

- A line where the Torque Fixture 2X Dowel pins sets zero in Z and is the Planar Alignment.

Beamline CSY (BEAMLINE$CSY)

- Use the BENCH$CSY setup except use the Torque Fixture plate to set zero in Y and the Spatial Alignment, transform +258.67 in Y to replicate Theoretical Beamline Center

- This coordinate system will tell us where the Wire Crosshairs are to a theoretical Beam Centerline.

- The wires should be offset ~ -0.400 in X, -0.200 in Y.

Statistical Data (mm.)

Out of Vacuum

10 Cycles (10 recordings)

Card Locations  
CSY = BENCH$CSY

| STATS | X-POS CARD | X-INTPT CARD | Y-INTPT CARD |
| --- | --- | --- | --- |
| AVG. | 0.03676 | -40.16612 | -95.40390 |
| Range | 0.02693 | 0.02422 | 0.00779 |
| St. Dev. | 0.00857 | 0.00726 | 0.00254 |

Wire Location

CSY = BEAMLINE$CSY

| STATS | X WIRE | Y WIRE |
| --- | --- | --- |
| AVG. | -0.39195 | -0.21148 |
| Range | 0.02694 | 0.00326 |
| St. Dev. | 0.00855 | 0.00097 |
| RMS Tol | 0.06500 | 0.03000 |
| Nom. Loc. | -0.40000 | -0.20000 |

Transformation WIRE$CSY to BFWTB$CSY

| Stats | X | Y | Z |
| --- | --- | --- | --- |
| AVG. | 94.71453 | -160.61536 | 35.89852 |
| Range | 0.00687 | 0.00347 | 0.05115 |
| St. Dev. | 0.00148 | 0.00098 | 0.01033 |

Under Vacuum

11 Cycles (11 recordings)

TB 5-8 Locations

CSY = BENCH$CSY

| TB 5 | X | Y | Z |
| --- | --- | --- | --- |
| AVG. | 94.15366 | 101.81124 | -37.51111 |
| Range | 0.01432 | 0.02706 | 0.02587 |
| St. Dev. | 0.00436 | 0.00894 | 0.00727 |
| RMS Tol | 0.06500 | 0.03000 | - |

Under Vacuum cont…

| TB 6 | X | Y | Z |
| --- | --- | --- | --- |
| AVG. | -94.71297 | 102.01512 | -37.38084 |
| Range | 0.01390 | 0.03022 | 0.01497 |
| St. Dev. | 0.00423 | 0.00968 | 0.00514 |
| RMS Tol | 0.06500 | 0.03000 | - |

| TB 7 | X | Y | Z |
| --- | --- | --- | --- |
| AVG. | -51.52804 | 101.83049 | 68.42183 |
| Range | 0.01010 | 0.01031 | 0.01292 |
| St. Dev. | 0.00289 | 0.00293 | 0.00454 |
| RMS Tol | 0.06500 | 0.03000 | - |

| TB 8 | X | Y | Z |
| --- | --- | --- | --- |
| AVG. | 51.05895 | 101.73264 | 68.36065 |
| Range | 0.00368 | 0.02439 | 0.02055 |
| St. Dev. | 0.00133 | 0.00776 | 0.00523 |
| RMS Tol | 0.06500 | 0.03000 | - |

Wire Location

CSY = BEAMLINE$CSY

| STATS | X WIRE | Y WIRE |
| --- | --- | --- |
| Vac. AVG. | -0.41198 | -0.18630 |
| No Vac. AVG. | -0.39195 | -0.21148 |
| Nom. Loc. | -0.40000 | -0.20000 |
| Range | 0.05891 | 0.00272 |
| St. Dev. | 0.01878 | 0.00084 |
| RMS Tol | 0.06500 | 0.03000 |

Transformation WIRE$CSY to BFWTB$CSY given to Ed Reese @ MMF

| Transformation | X | Y | Z |
| --- | --- | --- | --- |
| WIRE$CSY– BFWTB$CSY | 94.44864 | -160.08295 | 35.57212 |
| Transformation Matrix | 1.00000 | -0.00196 | -0.00101 |
|  | 0.00196 | 1.00000 | 0.00140 |
|  | -0.00101 | -0.00140 | 1.00000 |

Locations of ¼-20 Jo-Plugs Cylinders Intersected into Plane projected 12.7mm away from surface replicating a ½” (12.7mm) offset to center of Tooling Ball.

CSY = BEAMLINE$CSY

| INTPT | X | Y | Z |
| --- | --- | --- | --- |
| A | -26.93064 | 227.45522 | 68.97689 |
| B | 27.02135 | 227.48096 | 68.91648 |
| C | 79.75509 | 180.48270 | -68.27526 |
| D | -79.92870 | 180.50600 | -68.13640 |

Miscellaneous QC Checks

| Checklist | Y/N | Notes |
| --- | --- | --- |
| No Damage to BFW upon being Received? | Y |  |
| Prior to installation, ensure Wire Crosshairs are pertinent | Y |  |
| Wire card SN matches BFW SN | Y |  |
| Card has copper flex wires installed | Y |  |
| Connect Tabs from welded rod to Wire card | Y |  |
| No Thread damage on feedthru Threaded rods | N | Could not back out screws, used the ones Argonne supplied and there’s only 2 on each rod instead of 4. |
| Feedthru Welded Rods nuts tightly secured from 4 Ag nuts per weld rod | Y | See note above, only used 2 nuts on each rod. |
| Continuity from Top Feedthru to Card Connection Tabs | Y |  |
| Check 3 places for air fittings seal (RTV or Silicone/Teflon Tape) | Y |  |
| Connect Air @ 90psi, check for air leaks prior to actuation | Y |  |
| During actuation, ensure Limit Switches meet in Target In/Out Positions | Y |  |
| Ensure Light illuminates (LED) when turning on power supply | Y |  |
| Chamber: Look for aluminum coating on inside of chamber | Y |  |
| Chamber: Ensure flange clamp is made of Dichronite (Dark Gray) chain link/standard clamp | Y |  |
| Chamber: Ensure no loose hardware inside vacuum chamber | Y |  |
| Wire Crosshairs check prior to installation for vacuum | Y |  |
| BFW Vacuum Chamber is Pre Aligned (Girder Alignment) to X = 0.00 mm and Y = 258.67 mm to its Base | Y |  |
| Vacuum Leak Test: NIL (no indicated leak)? | Y |  |
| Wire Crosshairs are aligned to ±0.100 to X = -0.400 mm and Y = -0.200 mm | Y |  |
| Ready for Transportation Check: WC is target is “IN” position with 4 holding screws at the top. | Y |  |
| Installation of connector at the top of the Feedthru | Y |  |
| Check Continuity of the connector about the ceramic area | Y |  |
| Post Assy Check: Ensure plastic tubing is taped off at the top with mylar/electrical tape or top portion is closed for transportation (new metal shield protector) | Y |  |

# Text boxes

Y+

+Z

Z+

Y+

+X

+Y

X+

Transformation