LCLS BFW WIRECARD ASSY

FIDUCIAL REPORT

| Inspector: | Keith Caban |
| --- | --- |
| Responsible Engineer: | Venkat Srinivasan/Don Schafer |
| Date: | Thursday, May 22, 2008 |
| Work Order/Charge No.: | 23800-1 |
| Serial Number: | BFW – 09 / WC – 09 / BASE – 15 |
| URL | http://www-group/met/Quality/FIDUCIAL%20REPORTS/LCLS%20BFW/BFW-09_WC-09_BAS-15.pdf |

Coordinate System Setups

Bench CSY (BENCH$CSY)

- Plane measurement on top of Bench to set zero in Y and is the Spatial Alignment.

- Measure Pin and Diamond Pin to get centers, create a line between the two center points and this sets zero in Z and is the Planar Alignment.

- Take the symmetry point of the two pins this sets zero in X.

Bench Tooling Ball CSY (BTB$CSY)

- Create a plane using Tooling Balls 1, 2, & 3

  - This sets Zero in Y and is the spatial alignment

- Create an axis/line using TB’s 2 to 1 (+X)

  - This is the planar alignment

- TB 2 is a single point that sets zero in X, Y, Z.

BFW Tooling Ball CSY (BFWTB$CSY)

- Create a plane using Tooling Balls 5, 6, & 7

  - This sets Zero in Y and is the spatial alignment

- Create an axis/line using TB’s 6 to 5 (+X)

  - This is the planar alignment

- TB 6 is a single point that sets zero in X, Y, Z.

| View from Z+ or Downstream | View from X+ Side |
| --- | --- |

Wire CSY (WIRE$CSY)

- Back Plane (-Z) Side sets Zero in Z and is the Spatial Alignment

- Looking from +Z Side, Bottom of Card Sets Zero is the Planar Alignment

- The Intersection of the Left Side of the Card and the Bottom, sets Zero in X & Y

Wire offset from Edges per T. Nakashima’s Optical Measurement

| Wire | Wire Location | Wire Diameter |
| --- | --- | --- |
| Vertical Wire (X Location from Card) | 39.6092 | 0.0380 |
| Horizontal Wire (Y Location from Card) | 37.0105 | 0.0374 |

Support Stand CSY (BASE$CSY)

- A plane on Torque Fixture sets Zero in Y and is the Spatial Alignment

- The symmetry of the outer surfaces where “X” fine adjustments are sets zero in X.

- A line where the Torque Fixture 2X Dowel pins sets zero in Z and is the Planar Alignment.

Statistical Data

Out of Vacuum

Card Locations  
CSY = BENCH$CSY

| STATS | X-POS CARD | Y-INTPT CARD |
| --- | --- | --- |
| average | 0.07963 | -95.27622 |
| range | 0.03774 | 0.03243 |
| st dev | 0.00706 | 0.00697 |

Out of Vacuum Cont…

TB 5-8 Locations

CSY = BTB$CSY

| TB 5 | X | Y | Z |
| --- | --- | --- | --- |
| AVG. | 226.38705 | 89.58441 | -95.59554 |
| Range | 0.00870 | 0.02126 | 0.08643 |
| St. Dev. | 0.00164 | 0.00440 | 0.02180 |

| TB 6 | X | Y | Z |
| --- | --- | --- | --- |
| AVG. | 37.50763 | 89.68258 | -95.46461 |
| Range | 0.00844 | 0.11645 | 0.09285 |
| St. Dev. | 0.00175 | 0.02249 | 0.02215 |

| TB 7 | X | Y | Z |
| --- | --- | --- | --- |
| AVG. | 80.69758 | 89.66499 | 10.34147 |
| Range | 0.00375 | 0.10868 | 0.09096 |
| St. Dev. | 0.00081 | 0.02561 | 0.02199 |

| TB 8 | X | Y | Z |
| --- | --- | --- | --- |
| AVG. | 183.29051 | 89.59543 | 10.27191 |
| Range | 0.01027 | 0.09957 | 0.08709 |
| St. Dev. | 0.00376 | 0.02561 | 0.02160 |

Transformation BFWTB$CSY to BTB$CSY

| Stats | X | Y | Z | APA_PXPY | PA_PYPZ | PA_PXPZ |
| --- | --- | --- | --- | --- | --- | --- |
| AVG. | 37.50746 | 89.68818 | -95.46461 | -0.03148 | 89.99973 | 89.96028 |
| Range | 0.00844 | 0.02783 | 0.09285 | 0.01144 | 0.06486 | 0.00277 |
| St. Dev. | 0.00186 | 0.00564 | 0.02216 | 0.00218 | 0.01600 | 0.00059 |

Transformation WIRE$CSY to BTB$CSY

| Stats | X | Y | Z | APA_PXPY | PA_PYPZ | PA_PXPZ |
| --- | --- | --- | --- | --- | --- | --- |
| AVG. | 132.15877 | -70.97899 | -59.63082 | 0.05840 | 89.75070 | 89.99294 |
| Range | 0.03738 | 0.02880 | 0.27740 | 0.01102 | 0.06142 | 0.04398 |
| St. Dev. | 0.00702 | 0.00697 | 0.06874 | 0.00210 | 0.01428 | 0.00933 |

Transformation WIRE$CSY to BFWTB$CSY

| Stats | X | Y | Z | APA_PXPY | PA_PYPZ | PA_PXPZ |
| --- | --- | --- | --- | --- | --- | --- |
| AVG. | 94.71453 | -160.61536 | 35.89852 | 0.08987 | 89.75101 | 90.03282 |
| Range | 0.00687 | 0.00347 | 0.05115 | 0.00304 | 0.00903 | 0.04439 |
| St. Dev. | 0.00148 | 0.00098 | 0.01033 | 0.00069 | 0.00266 | 0.00937 |

Under Vacuum

**CMM CRASHED DURING THE 32nd CYCLE, machine had to re-zero and could not replicate the BENCH$CSY w/o the risk of losing another gasket and time/effort so there will be 2 sets of statistics for “UNDER VACUUM” data

Pre-Crash Data: results thru 22 cycles & 8 data recordings

TB 5-8 Locations

CSY = BENCH$CSY

| TB 5 | X | Y | Z |
| --- | --- | --- | --- |
| AVG. | 93.91535 | 102.37560 | -38.23587 |
| Range | 0.00162 | 0.01154 | 0.01005 |
| St. Dev. | 0.00056 | 0.00419 | 0.00406 |

| TB 6 | X | Y | Z |
| --- | --- | --- | --- |
| AVG. | -94.96640 | 102.40471 | -37.95236 |
| Range | 0.00153 | 0.00994 | 0.05520 |
| St. Dev. | 0.00053 | 0.00361 | 0.02231 |

| TB 7 | X | Y | Z |
| --- | --- | --- | --- |
| AVG. | -51.68977 | 102.46170 | 67.82019 |
| Range | 0.03598 | 0.01009 | 0.04208 |
| St. Dev. | 0.01419 | 0.00405 | 0.01660 |

| TB 8 | X | Y | Z |
| --- | --- | --- | --- |
| AVG. | 50.90679 | 102.43152 | 67.66268 |
| Range | 0.03585 | 0.00205 | 0.00683 |
| St. Dev. | 0.01419 | 0.00072 | 0.00240 |

Transformation BFWTB$CSY to BASE$CSY

| Stats | X | Y | Z | APA_PXPY | PA_PYPZ | PA_PXPZ |
| --- | --- | --- | --- | --- | --- | --- |
| AVG. | -95.82958 | 419.08004 | -101.38326 | -0.00249 | 89.98115 | 89.64762 |
| Range | 0.00128 | 0.00995 | 0.05520 | 0.00609 | 0.00352 | 0.01979 |
| St. Dev. | 0.00045 | 0.00361 | 0.02231 | 0.00235 | 0.00128 | 0.00799 |

Transformation WIRE$CSY to BENCH$CSY

| Stats | X | Y | Z | APA_PXPY | PA_PYPZ | PA_PXPZ |
| --- | --- | --- | --- | --- | --- | --- |
| AVG. | -0.22306 | -58.20336 | -2.09791 | 0.08126 | 89.71808 | 89.94299 |
| Range | 0.02886 | 0.00179 | 0.01545 | 0.00609 | 0.00353 | 0.01976 |
| St. Dev. | 0.01106 | 0.00056 | 0.00565 | 0.00234 | 0.00128 | 0.00798 |

Pre-Crash Data: results thru 22-100 cycles & 13 data recordings

TB 5-8 Locations

CSY = BENCH$CSY

| TB 5 | X | Y | Z |
| --- | --- | --- | --- |
| AVG. | 93.92803 | 102.46187 | -38.17051 |
| Range | 0.00193 | 0.01049 | 0.01123 |
| St. Dev. | 0.00060 | 0.00273 | 0.00364 |

| TB 6 | X | Y | Z |
| --- | --- | --- | --- |
| AVG. | -94.95426 | 102.48295 | -37.87777 |
| Range | 0.00190 | 0.01220 | 0.06762 |
| St. Dev. | 0.00063 | 0.00337 | 0.02008 |

| TB 7 | X | Y | Z |
| --- | --- | --- | --- |
| AVG. | -51.66416 | 102.54175 | 67.89266 |
| Range | 0.09591 | 0.01348 | 0.04651 |
| St. Dev. | 0.02478 | 0.00367 | 0.01431 |

| TB 8 | X | Y | Z |
| --- | --- | --- | --- |
| AVG. | 50.92450 | 102.51559 | 67.72996 |
| Range | 0.04342 | 0.00260 | 0.00649 |
| St. Dev. | 0.01298 | 0.00073 | 0.00176 |

Transformation BFWTB$CSY to BASE$CSY

| Stats | X | Y | Z | APA_PXPY | PA_PYPZ | PA_PXPZ |
| --- | --- | --- | --- | --- | --- | --- |
| AVG. | -95.81710 | 419.15825 | -101.30870 | -0.00029 | 89.98115 | 89.64378 |
| Range | 0.00201 | 0.01221 | 0.06761 | 0.00556 | 0.00350 | 0.02353 |
| St. Dev. | 0.00062 | 0.00337 | 0.02008 | 0.00170 | 0.00098 | 0.00687 |

Transformation WIRE$CSY to BENCH$CSY

| Stats | X | Y | Z | APA_PXPY | PA_PYPZ | PA_PXPZ |
| --- | --- | --- | --- | --- | --- | --- |
| AVG. | -0.20234 | -58.12110 | -2.02798 | 0.08370 | 89.71809 | 89.94020 |
| Range | 0.02966 | 0.00560 | 0.02096 | 0.00555 | 0.00351 | 0.02351 |
| St. Dev. | 0.00894 | 0.00135 | 0.00584 | 0.00172 | 0.00098 | 0.00716 |

Miscellaneous QC Checks

| Checklist | Y/N | Notes |
| --- | --- | --- |
| Out of Crate Damage | N | Just has Pink Foam all over. |
| Wire card SN matches BFW SN | Y |  |
| Card has copper flex wires installed | Y |  |
| Connect Tabs from welded rod to Wire card | Y |  |
| Feedthru Welded Rods nuts tightly secured from 4 Ag nuts per weld rod | Y | Argonne only provided 1 nut on each weld rod, removed and added 4 on each |
| Continuity from Top Feedthru to Card Connection Tabs | Y |  |
| Check 3 places for air fittings seal (RTV or Silicone/Teflon Tape) | Y |  |
| Connect Air @ 90psi, check for air leaks prior to actuation | Y |  |
| During actuation, ensure Limit Switches meet in Target In/Out Positions | Y | Noted new limit switch |
| Ensure Light illuminates (LED) when turning on power supply | Y |  |
| Chamber: Look for aluminum coating on inside of chamber | Y |  |
| Chamber: Ensure flange clamp is made of Dichronite (Dark Gray) chain link/standard clamp | Y |  |
| Chamber: Ensure no loose hardware inside vacuum chamber | N | Found 2 8-32 hex nuts, 1 inside chamber, 1 by standard clamp. |
| Vacuum Leak Test: NIL (no indicated leak)? | Y |  |
| Ready for Transportation Check: WC is target is “IN” position with 4 holding screws at the top. | Y |  |
| Post Assy Check: Ensure plastic tubing is taped off at the top with mylar/electrical tape | Y |  |

Bottom of Form

# Text boxes

Y+

+Z

Z+

Y+

+X

+Y

X+

Transformation