LCLS BFW WIRECARD ASSY

INSPECTION/FIDUCIAL REPORT

| Inspector: | Keith Caban |
| --- | --- |
| Responsible Engineer: | Venkat Srinivasan |
| Date: | Tuesday, June 17, 2008 |
| Work Order/Charge No.: | 23800-1 |
| Serial Number: | BFW – 13 / VC – 13 / WC – 13 / BASE – 13 |
| Units of Measure: | Metric – Millimeters (mm.) |
| URL | http://www-group/met/Quality/FIDUCIAL%20REPORTS/LCLS%20BFW/BFW-13_VC-13_WC-13_BAS-13.pdf |

Coordinate System Setups

Bench CSY (BENCH$CSY)

(The only CSY in Both States of measurement - out and under vacuum)

- Plane measurement on top of Bench to set zero in Y and is the Spatial Alignment.

- Measure Pin and Diamond Pin to get centers, create a line between the two center points and this sets zero in Z and is the Planar Alignment.

- Take the symmetry point of the two pins this sets zero in X.

Bench Tooling Ball CSY (BTB$CSY)

- Create a plane using Tooling Balls 1, 2, & 3

  - This sets Zero in Y and is the spatial alignment

- Create an axis/line using TB’s 2 to 1 (+X)

  - This is the planar alignment

- TB 2 is a single point that sets zero in X, Y, Z.

BFW Tooling Ball CSY (BFWTB$CSY)

- Create a plane using Tooling Balls 5, 6, & 7

  - This sets Zero in Y and is the spatial alignment

- Create an axis/line using TB’s 6 to 5 (+X)

  - This is the planar alignment

- TB 6 is a single point that sets zero in X, Y, Z.

| View from Z+ or Downstream | View from X+ Side |
| --- | --- |

Wire CSY (WIRE$CSY)

- Back Plane (-Z) Side sets Zero in Z and is the Spatial Alignment

- Looking from +Z Side, Bottom of Card Sets Zero is the Planar Alignment

- The Intersection of the Left Side of the Card and the Bottom, sets Zero in X & Y

Wire offset from Edges per T. Nakashima’s Optical Measurement

| Wire | Wire Location | Wire Diameter |
| --- | --- | --- |
| Vertical Wire (X Location from Card) | 39.6044 | 0.0387 |
| Horizontal Wire (Y Location from Card) | 37.0070 | 0.0402 |

Support Stand CSY (BASE$CSY)

- A plane on Torque Fixture sets Zero in Y and is the Spatial Alignment

- The symmetry of the outer surfaces where “X” fine adjustments are sets zero in X.

- A line where the Torque Fixture 2X Dowel pins sets zero in Z and is the Planar Alignment.

Beamline CSY (BEAMLINE$CSY)

- Use the BENCH$CSY setup except use the Torque Fixture plate to set zero in Y and the Spatial Alignment, transform +258.67 in Y to replicate Theoretical Beamline Center

- This coordinate system will tell us where the Wire Crosshairs are to a theoretical Beam Centerline.

- The wires should be offset ~ -0.400 in X, -0.200 in Y.

Statistical Data (mm.)

Out of Vacuum

10 Cycles (10 recordings)

Card Locations  
CSY = BENCH$CSY

| STATS | X-POS CARD | X-INTPT CARD | Y-INTPT CARD |
| --- | --- | --- | --- |
| AVG. | -0.00764 | -40.19406 | -94.90638 |
| Range | 0.00347 | 0.00373 | 0.00097 |
| St. Dev. | 0.00113 | 0.00127 | 0.00028 |

Wire Location

CSY = BEAMLINE$CSY

| Stats | X | Y |
| --- | --- | --- |
| AVG. | -0.43161 | -0.08121** |
| Range | 0.00335 | 0.00056 |
| St. Dev. | 0.00110 | 0.00022 |
| Nom. | -0.40000 | -0.20000 |
| RMS Tol | 0.06500 | 0.03000 |

**Did not make “Y” adjustment outside of vacuum.

Transformation WIRE$CSY TO BENCH$CSY

| Stats | X | Y | Z | APA_PXPY | PA_PYPZ | PA_PXPZ |
| --- | --- | --- | --- | --- | --- | --- |
| AVG. | -0.43161 | -58.08121 | -2.74982 | -0.24985 | 89.86261 | 89.81926 |
| Range | 0.00335 | 0.00056 | 0.00296 | 0.00097 | 0.00069 | 0.00103 |
| St. Dev. | 0.00110 | 0.00022 | 0.00091 | 0.00036 | 0.00026 | 0.00030 |
| RMS Tol | 0.06500 | 0.03000 | - | - | - | - |

Transformation WIRE$CSY to BFWTB$CSY

| Stats | X | Y | Z | APA_PXPY | PA_PYPZ | PA_PXPZ |
| --- | --- | --- | --- | --- | --- | --- |
| AVG. | 94.48174 | -160.34396 | 35.90405 | -0.10883 | 89.73230 | 90.06517 |
| Range | 0.00115 | 0.00052 | 0.00243 | 0.00046 | 0.00104 | 0.00046 |
| St. Dev. | 0.00038 | 0.00017 | 0.00075 | 0.00016 | 0.00031 | 0.00014 |
| RMS Tol | 0.06500 | 0.03000 | - | - | - | - |

Under Vacuum

10 Cycles (10 recordings)

TB 5-8 Locations

CSY = BENCH$CSY

| TB 5 | X | Y | Z |
| --- | --- | --- | --- |
| AVG. | 94.30514 | 101.81864 | -38.58959 |
| Range | 0.01852 | 0.00597 | 0.02107 |
| St. Dev. | 0.00683 | 0.00209 | 0.00831 |
| RMS Tol | 0.06500 | 0.03000 | - |

| TB 6 | X | Y | Z |
| --- | --- | --- | --- |
| AVG. | -94.57432 | 102.40375 | -37.94866 |
| Range | 0.01847 | 0.00651 | 0.02375 |
| St. Dev. | 0.00683 | 0.00200 | 0.00886 |
| RMS Tol | 0.06500 | 0.03000 | - |

| TB 7 | X | Y | Z |
| --- | --- | --- | --- |
| AVG. | -51.08183 | 102.28872 | 67.75400 |
| Range | 0.00827 | 0.00138 | 0.01328 |
| St. Dev. | 0.00295 | 0.00045 | 0.00496 |
| RMS Tol | 0.06500 | 0.03000 | - |

| TB 8 | X | Y | Z |
| --- | --- | --- | --- |
| AVG. | 51.49276 | 101.97088 | 67.40274 |
| Range | 0.00759 | 0.00570 | 0.01101 |
| St. Dev. | 0.00263 | 0.00182 | 0.00390 |
| RMS Tol | 0.06500 | 0.03000 | - |

Wire Location

CSY = BEAMLINE$CSY

| Stats | X | Y |
| --- | --- | --- |
| AVG. | -0.46800 | -0.20406 |
| Range | 0.01071 | 0.00112 |
| St. Dev. | 0.00340 | 0.00038 |
| Nom. | -0.40000 | -0.20000 |
| RMS Tol | 0.06500 | 0.03000 |

Transformation WIRE$CSY TO BENCH$CSY

| Stats | X | Y | Z | APA_PXPY | PA_PYPZ | PA_PXPZ |
| --- | --- | --- | --- | --- | --- | --- |
| AVG. | -0.46856 | -58.22569 | -2.33584 | -0.28661 | 89.72204 | 89.87010 |
| Range | 0.01071 | 0.00112 | 0.00581 | 0.00375 | 0.00178 | 0.01343 |
| St. Dev. | 0.00368 | 0.00040 | 0.00182 | 0.00120 | 0.00055 | 0.00520 |
| RMS Tol | 0.06500 | 0.03000 | - | - | - | - |

Locations of ¼-20 Jo-Plugs Cylinders Intersected into Plane projected 12.7mm away from surface replicating a ½” (12.7mm) offset to center of Tooling Ball.

CSY = BEAMLINE$CSY

| INTPT | X | Y | Z |
| --- | --- | --- | --- |
| 1 | -26.98548 | 227.65687 | 68.40414 |
| 2 | 26.98643 | 227.64760 | 68.41702 |
| 3 | 79.87051 | 180.23185 | -68.58461 |
| 4 | -79.80670 | 180.24252 | -68.62311 |

# Text boxes

Y+

+Z

Z+

Y+

+X

+Y

X+

Transformation