LCLS BFW WIRECARD ASSY

INSPECTION/FIDUCIAL REPORT

| Inspector: | Keith Caban |
| --- | --- |
| Responsible Engineer: | Don Schafer |
| Date: | Thursday, July 24, 2008 |
| Work Order/Charge No.: | 23800-1 |
| Serial Number: | BFW – 14 / VC – 14 / WC – 14 / BASE – 32 |
| Units of Measure: | Metric – Millimeters (mm.) |
| URL | http://www-group/met/Quality/FIDUCIAL%20REPORTS/LCLS%20BFW/BFW-14_VC-14_WC-14_BAS-32.pdf |

Coordinate System Setups

Bench CSY (BENCH$CSY)

(The only CSY in Both States of measurement - out and under vacuum)

- Plane measurement on top of Bench to set zero in Y and is the Spatial Alignment.

- Measure Pin and Diamond Pin to get centers, create a line between the two center points and this sets zero in Z and is the Planar Alignment.

- Take the symmetry point of the two pins this sets zero in X.

Bench Tooling Ball CSY (BTB$CSY)

- Create a plane using Tooling Balls 1, 2, & 3

  - This sets Zero in Y and is the spatial alignment

- Create an axis/line using TB’s 2 to 1 (+X)

  - This is the planar alignment

- TB 2 is a single point that sets zero in X, Y, Z.

BFW Tooling Ball CSY (BFWTB$CSY)

- Create a plane using Tooling Balls 5, 6, & 7

  - This sets Zero in Y and is the spatial alignment

- Create an axis/line using TB’s 6 to 5 (+X)

  - This is the planar alignment

- TB 6 is a single point that sets zero in X, Y, Z.

| View from Z+ or Downstream | View from X+ Side |
| --- | --- |

Wire CSY (WIRE$CSY)

- Back Plane (-Z) Side sets Zero in Z and is the Spatial Alignment

- Looking from +Z Side, Bottom of Card Sets Zero is the Planar Alignment

- The Intersection of the Left Side of the Card and the Bottom, sets Zero in X & Y

Wire offset from Edges per T. Nakashima’s Optical Measurement

| Wire | Wire Location | Wire Diameter |
| --- | --- | --- |
| Vertical Wire (X Location from Card) | 39.6007 | 0.0374 |
| Horizontal Wire (Y Location from Card) | 37.0050 | 0.0397 |

Support Stand CSY (BASE$CSY)

- A plane on Torque Fixture sets Zero in Y and is the Spatial Alignment

- The symmetry of the outer surfaces where “X” fine adjustments are sets zero in X.

- A line where the Torque Fixture 2X Dowel pins sets zero in Z and is the Planar Alignment.

Beamline CSY (BEAMLINE$CSY)

- Use the BENCH$CSY setup except use the Torque Fixture plate to set zero in Y and the Spatial Alignment, transform +258.67 in Y to replicate Theoretical Beamline Center

- This coordinate system will tell us where the Wire Crosshairs are to a theoretical Beam Centerline.

- The wires should be offset ~ -0.400 in X, -0.200 in Y.

Statistical Data (mm.)

Out of Vacuum

10 Cycles (10 recordings)

Card Locations  
CSY = BENCH$CSY

| STATS | X-POS CARD | X-INTPT CARD | Y-INTPT CARD |
| --- | --- | --- | --- |
| 0.04187 | 0.03259 | -39.87818 | -95.34736 |
| 0.00150 | 0.13038 | 0.15463 | 0.03735 |
| 0.00056 | 0.03542 | 0.04203 | 0.01015 |

Wire Location

CSY = BEAMLINE$CSY

| Stats | X | Y |
| --- | --- | --- |
| AVG. | -0.38838 | -0.23228 |
| Range | 0.13230 | 0.01399 |
| St. Dev. | 0.03589 | 0.00372 |
| Nom. | -0.40000 | -0.20000 |
| RMS Tol | 0.06500 | 0.03000 |

Transformation WIRE$CSY TO BENCH$CSY

| Stats | X | Y | Z | APA_PXPY | PA_PYPZ | PA_PXPZ |
| --- | --- | --- | --- | --- | --- | --- |
| AVG. | -0.38838 | -58.23228 | -3.14271 | 0.17132 | 90.01502 | 90.15364 |
| Range | 0.13230 | 0.01399 | 0.01108 | 0.03418 | 0.00207 | 0.00867 |
| St. Dev. | 0.03589 | 0.00372 | 0.00286 | 0.00942 | 0.00069 | 0.00253 |
| RMS Tol | 0.06500 | 0.03000 | - | - | - | - |

Transformation WIRE$CSY to BFWTB$CSY

| Stats | X | Y | Z | APA_PXPY | PA_PYPZ | PA_PXPZ |
| --- | --- | --- | --- | --- | --- | --- |
| AVG. | 94.00301 | -160.59202 | 35.35521 | 0.09503 | 90.02438 | 90.02645 |
| Range | 0.00350 | 0.00057 | 0.00375 | 0.00131 | 0.00195 | 0.00493 |
| St. Dev. | 0.00099 | 0.00020 | 0.00131 | 0.00038 | 0.00066 | 0.00158 |
| RMS Tol | 0.06500 | 0.03000 | - | - | - | - |

Under Vacuum

10 Cycles (10 recordings)

TB 5-8 Locations

CSY = BENCH$CSY

| TB 5 | X | Y | Z |
| --- | --- | --- | --- |
| AVG. | 94.25653 | 102.41377 | -38.30901 |
| Range | 0.00901 | 0.00877 | 0.00579 |
| St. Dev. | 0.00302 | 0.00276 | 0.00187 |
| RMS Tol | 0.06500 | 0.03000 | - |

| TB 6 | X | Y | Z |
| --- | --- | --- | --- |
| AVG. | -94.61386 | 102.10482 | -38.46394 |
| Range | 0.00945 | 0.01400 | 0.00329 |
| St. Dev. | 0.00318 | 0.00518 | 0.00121 |
| RMS Tol | 0.06500 | 0.03000 | - |

| TB 7 | X | Y | Z |
| --- | --- | --- | --- |
| AVG. | -51.58489 | 102.61350 | 67.41970 |
| Range | 0.00587 | 0.00976 | 0.00151 |
| St. Dev. | 0.00208 | 0.00356 | 0.00053 |
| RMS Tol | 0.06500 | 0.03000 | - |

| TB 8 | X | Y | Z |
| --- | --- | --- | --- |
| AVG. | 51.00382 | 102.79229 | 67.49809 |
| Range | 0.00589 | 0.00296 | 0.00352 |
| St. Dev. | 0.00203 | 0.00088 | 0.00119 |
| RMS Tol | 0.06500 | 0.03000 | - |

Wire Location

CSY = BEAMLINE$CSY

| Stats | X | Y |
| --- | --- | --- |
| AVG. | -0.36626 | -0.23220 |
| Range | 0.02768 | 0.00345 |
| St. Dev. | 0.00939 | 0.00134 |
| Nom. | -0.40000 | -0.20000 |
| RMS Tol | 0.06500 | 0.03000 |

Transformation WIRE$CSY TO BENCH$CSY

| Stats | X | Y | Z | APA_PXPY | PA_PYPZ | PA_PXPZ |
| --- | --- | --- | --- | --- | --- | --- |
| AVG. | -0.37863 | -58.18541 | -2.36793 | 0.18893 | 89.78756 | 90.07309 |
| Range | 0.02767 | 0.00345 | 0.00215 | 0.00691 | 0.00060 | 0.00277 |
| St. Dev. | 0.00939 | 0.00134 | 0.00066 | 0.00240 | 0.00022 | 0.00090 |
| RMS Tol | 0.06500 | 0.03000 | - | - | - | - |

Locations of ¼-20 Jo-Plugs Cylinders Intersected into Plane projected 12.7mm away from surface replicating a ½” (12.7mm) offset to center of Tooling Ball.

CSY = BEAMLINE$CSY

| INTPT | X | Y | Z |
| --- | --- | --- | --- |
| 1 | -27.09029 | 227.65260 | 68.23698 |
| 2 | 26.88275 | 227.66006 | 68.22988 |
| 3 | 79.72081 | 180.17599 | -68.73940 |
| 4 | -79.90053 | 180.12181 | -68.71129 |

Wire Positioning/Locations with respect to the BEAMLINE$CSY during processes:

| Wire Position | X | Y |
| --- | --- | --- |
| Avg. Bench Pos. | -0.38838 | -0.23228 |
| After Assy. On Vac. Chamber (actuated) | -0.43938 | -0.35928 |
| After Pumpdown held in Up Pos. | -0.61914 | -0.34668 |
| 1 Cycle in Vac | -0.43752 | -0.36472 |
| Avg. Vac. Pos. after Y adjust | -0.36626 | -0.23220 |
| After Vent still in held in Up Pos. | -0.57363 | -0.21013 |
| Actuated after Vent | -0.33049 | -0.22852 |

# Text boxes

Y+

+Z

Z+

Y+

+X

+Y

X+

Transformation