LCLS BFW WIRECARD ASSY

INSPECTION/FIDUCIAL REPORT

| Inspector: | Keith Caban |
| --- | --- |
| Responsible Engineer: | Venkat Srinivasan |
| Date: | Thursday, July 03, 2008 |
| Work Order/Charge No.: | 23800-1 |
| Serial Number: | BFW – 16 / VC – 16 / WC – 16 / BASE – 18 |
| Units of Measure: | Metric – Millimeters (mm.) |
| URL | http://www-group/met/Quality/FIDUCIAL%20REPORTS/LCLS%20BFW/BFW-16_VC-16_WC-16_BAS-18.pdf |

Coordinate System Setups

Bench CSY (BENCH$CSY)

(The only CSY in Both States of measurement - out and under vacuum)

- Plane measurement on top of Bench to set zero in Y and is the Spatial Alignment.

- Measure Pin and Diamond Pin to get centers, create a line between the two center points and this sets zero in Z and is the Planar Alignment.

- Take the symmetry point of the two pins this sets zero in X.

Bench Tooling Ball CSY (BTB$CSY)

- Create a plane using Tooling Balls 1, 2, & 3

  - This sets Zero in Y and is the spatial alignment

- Create an axis/line using TB’s 2 to 1 (+X)

  - This is the planar alignment

- TB 2 is a single point that sets zero in X, Y, Z.

BFW Tooling Ball CSY (BFWTB$CSY)

- Create a plane using Tooling Balls 5, 6, & 7

  - This sets Zero in Y and is the spatial alignment

- Create an axis/line using TB’s 6 to 5 (+X)

  - This is the planar alignment

- TB 6 is a single point that sets zero in X, Y, Z.

| View from Z+ or Downstream | View from X+ Side |
| --- | --- |

Wire CSY (WIRE$CSY)

- Back Plane (-Z) Side sets Zero in Z and is the Spatial Alignment

- Looking from +Z Side, Bottom of Card Sets Zero is the Planar Alignment

- The Intersection of the Left Side of the Card and the Bottom, sets Zero in X & Y

Wire offset from Edges per T. Nakashima’s Optical Measurement

| Wire | Wire Location | Wire Diameter |
| --- | --- | --- |
| Vertical Wire (X Location from Card) | 39.6075 | 0.0363 |
| Horizontal Wire (Y Location from Card) | 37.0017 | 0.0392 |

Support Stand CSY (BASE$CSY)

- A plane on Torque Fixture sets Zero in Y and is the Spatial Alignment

- The symmetry of the outer surfaces where “X” fine adjustments are sets zero in X.

- A line where the Torque Fixture 2X Dowel pins sets zero in Z and is the Planar Alignment.

Beamline CSY (BEAMLINE$CSY)

- Use the BENCH$CSY setup except use the Torque Fixture plate to set zero in Y and the Spatial Alignment, transform +258.67 in Y to replicate Theoretical Beamline Center

- This coordinate system will tell us where the Wire Crosshairs are to a theoretical Beam Centerline.

- The wires should be offset ~ -0.400 in X, -0.200 in Y.

Statistical Data (mm.)

Out of Vacuum

10 Cycles (10 recordings)

Card Locations  
CSY = BENCH$CSY

| STATS | X-POS CARD | X-INTPT CARD | Y-INTPT CARD |
| --- | --- | --- | --- |
| AVG. | 0.03618 | -40.75043 | -95.44442 |
| Range | 0.04412 | 0.14916 | 0.01012 |
| St. Dev. | 0.01380 | 0.05168 | 0.00330 |

Wire Location

CSY = BEAMLINE$CSY

| Stats | X | Y |
| --- | --- | --- |
| AVG. | -0.38708 | -0.21622 |
| Range | 0.04419 | 0.00165 |
| St. Dev. | 0.01383 | 0.00066 |
| Nom. | -0.40000 | -0.20000 |
| RMS Tol | 0.06500 | 0.03000 |

Transformation WIRE$CSY TO BENCH$CSY

| Stats | X | Y | Z | APA_PXPY | PA_PYPZ | PA_PXPZ |
| --- | --- | --- | --- | --- | --- | --- |
| AVG. | -0.38708 | -58.21622 | -3.27770 | 0.33838 | 90.22179 | 90.00272 |
| Range | 0.04419 | 0.00165 | 0.04419 | 0.01330 | 0.01141 | 0.00534 |
| St. Dev. | 0.01383 | 0.00066 | 0.01376 | 0.00413 | 0.00356 | 0.00165 |
| RMS Tol | 0.06500 | 0.03000 | - | - | - | - |

Transformation WIRE$CSY to BFWTB$CSY

| Stats | X | Y | Z | APA_PXPY | PA_PYPZ | PA_PXPZ |
| --- | --- | --- | --- | --- | --- | --- |
| AVG. | 93.37386 | -160.84725 | 35.65308 | 0.02929 | 89.91594 | 89.98503 |
| Range | 0.00594 | 0.00272 | 0.00441 | 0.00196 | 0.00169 | 0.00185 |
| St. Dev. | 0.00178 | 0.00094 | 0.00122 | 0.00065 | 0.00047 | 0.00062 |
| RMS Tol | 0.06500 | 0.03000 | - | - | - | - |

Under Vacuum

10 Cycles (10 recordings)

TB 5-8 Locations

CSY = BENCH$CSY

| TB 5 | X | Y | Z |
| --- | --- | --- | --- |
| AVG. | 94.21389 | 103.14434 | -37.76802 |
| Range | 0.00737 | 0.00688 | 0.00820 |
| St. Dev. | 0.00234 | 0.00258 | 0.00261 |
| RMS Tol | 0.06500 | 0.03000 | - |

| TB 6 | X | Y | Z |
| --- | --- | --- | --- |
| AVG. | -94.64972 | 102.20924 | -37.52269 |
| Range | 0.00758 | 0.00345 | 0.01027 |
| St. Dev. | 0.00245 | 0.00116 | 0.00326 |
| RMS Tol | 0.06500 | 0.03000 | - |

| TB 7 | X | Y | Z |
| --- | --- | --- | --- |
| AVG. | -51.39675 | 102.06872 | 68.26850 |
| Range | 0.00326 | 0.00261 | 0.00566 |
| St. Dev. | 0.00103 | 0.00086 | 0.00181 |
| RMS Tol | 0.06500 | 0.03000 | - |

| TB 8 | X | Y | Z |
| --- | --- | --- | --- |
| AVG. | 51.18998 | 102.57942 | 68.11415 |
| Range | 0.00353 | 0.00397 | 0.00421 |
| St. Dev. | 0.00105 | 0.00152 | 0.00139 |
| RMS Tol | 0.06500 | 0.03000 | - |

Wire Location

CSY = BEAMLINE$CSY

| Stats | X | Y |
| --- | --- | --- |
| AVG. | -0.42515 | -0.27334 |
| Range | 0.00764 | 0.00220 |
| St. Dev. | 0.00269 | 0.00067 |
| Nom. | -0.40000 | -0.20000 |
| RMS Tol | 0.06500 | 0.03000 |

Transformation WIRE$CSY TO BENCH$CSY

| Stats | X | Y | Z | APA_PXPY | PA_PYPZ | PA_PXPZ |
| --- | --- | --- | --- | --- | --- | --- |
| AVG. | -0.43319 | -58.29120 | -2.53132 | 0.31350 | 90.10788 | 89.90961 |
| Range | 0.00764 | 0.00220 | 0.00313 | 0.00300 | 0.00122 | 0.00561 |
| St. Dev. | 0.00270 | 0.00075 | 0.00103 | 0.00109 | 0.00045 | 0.00177 |
| RMS Tol | 0.06500 | 0.03000 | - | - | - | - |

Locations of ¼-20 Jo-Plugs Cylinders Intersected into Plane projected 12.7mm away from surface replicating a ½” (12.7mm) offset to center of Tooling Ball.

CSY = BEAMLINE$CSY

| INTPT | X | Y | Z |
| --- | --- | --- | --- |
| 1 | -27.20282 | 227.55292 | 68.62271 |
| 2 | 26.83392 | 227.58422 | 68.57984 |
| 3 | 79.54441 | 180.36487 | -68.54096 |
| 4 | -80.09338 | 180.31729 | -68.41221 |

# Text boxes

Y+

+Z

Z+

Y+

+X

+Y

X+

Transformation