LCLS BFW WIRECARD ASSY

INSPECTION/FIDUCIAL REPORT

| Inspector: | Keith Caban |
| --- | --- |
| Responsible Engineer: | Don Schafer |
| Date: | Wednesday, September 24, 2008 |
| Work Order/Charge No.: | 23800-1 |
| Serial Number: | BFW – 21 / VC – 21 / WC – 21 / BASE – 26 |
| Units of Measure: | Metric – Millimeters (mm.) |
| URL | http://www-group/met/Quality/FIDUCIAL%20REPORTS/LCLS%20BFW/BFW-21_VC-21_WC-21_BAS-26.pdf |

Coordinate System Setups

Bench CSY (BENCH$CSY)

(The only CSY in Both States of measurement - out and under vacuum)

- Plane measurement on top of Bench to set zero in Y and is the Spatial Alignment.

- Measure Pin and Diamond Pin to get centers, create a line between the two center points and this sets zero in Z and is the Planar Alignment.

- Take the symmetry point of the two pins this sets zero in X.

Bench Tooling Ball CSY (BTB$CSY)

- Create a plane using Tooling Balls 1, 2, & 3

  - This sets Zero in Y and is the spatial alignment

- Create an axis/line using TB’s 2 to 1 (+X)

  - This is the planar alignment

- TB 2 is a single point that sets zero in X, Y, Z.

BFW Tooling Ball CSY (BFWTB$CSY)

- Create a plane using Tooling Balls 5, 6, & 7

  - This sets Zero in Y and is the spatial alignment

- Create an axis/line using TB’s 6 to 5 (+X)

  - This is the planar alignment

- TB 6 is a single point that sets zero in X, Y, Z.

| View from Z+ or Downstream | View from X+ Side |
| --- | --- |

Wire CSY (WIRE$CSY)

- Back Plane (-Z) Side sets Zero in Z and is the Spatial Alignment

- Looking from +Z Side, Bottom of Card Sets Zero is the Planar Alignment

- The Intersection of the Left Side of the Card and the Bottom, sets Zero in X & Y

Wire offset from Edges per T. Nakashima’s Optical Measurement

| Wire | Wire Location | Wire Diameter |
| --- | --- | --- |
| Vertical Wire (X Location from Card) | 39.6090 | 0.0383 |
| Horizontal Wire (Y Location from Card) | 37.0029 | 0.0376 |

Support Stand CSY (BASE$CSY)

- A plane on Torque Fixture sets Zero in Y and is the Spatial Alignment

- The symmetry of the outer surfaces where “X” fine adjustments are sets zero in X.

- A line where the Torque Fixture 2X Dowel pins sets zero in Z and is the Planar Alignment.

Beamline CSY (BEAMLINE$CSY)

- Use the BENCH$CSY setup except use the Torque Fixture plate to set zero in Y and the Spatial Alignment, transform +258.67 in Y to replicate Theoretical Beamline Center

- This coordinate system will tell us where the Wire Crosshairs are to a theoretical Beam Centerline.

- The wires should be offset ~ -0.400 in X, -0.200 in Y.

Statistical Data (mm.)

Out of Vacuum

10 Cycles (10 recordings)

Card Locations  
CSY = BENCH$CSY

| STATS | X-POS CARD | X-INTPT CARD | Y-INTPT CARD |
| --- | --- | --- | --- |
| AVG. | 0.01158 | -40.10252 | -95.04465 |
| Range | 0.03215 | 0.03871 | 0.00602 |
| St. Dev. | 0.00991 | 0.01192 | 0.00177 |

Wire Location

CSY = BEAMLINE$CSY

| Stats | X | Y |
| --- | --- | --- |
| AVG. | -0.40981 | -0.14183 |
| Range | 0.03203 | 0.00158 |
| St. Dev. | 0.00984 | 0.00056 |
| Nom. | -0.40000 | -0.20000 |
| RMS Tol | 0.06500 | 0.03000 |

Transformation WIRE$CSY TO BENCH$CSY

| Stats | X | Y | Z | APA_PXPY | PA_PYPZ | PA_PXPZ |
| --- | --- | --- | --- | --- | --- | --- |
| AVG. | -0.40981 | -58.14183 | -2.78858 | -0.13142 | 89.72798 | 89.89896 |
| Range | 0.03203 | 0.00158 | 0.03405 | 0.00986 | 0.01068 | 0.00149 |
| St. Dev. | 0.00984 | 0.00056 | 0.01054 | 0.00299 | 0.00333 | 0.00045 |
| RMS Tol | 0.06500 | 0.03000 | - | - | - | - |

Transformation WIRE$CSY to BFWTB$CSY

| Stats | X | Y | Z | APA_PXPY | PA_PYPZ | PA_PXPZ |
| --- | --- | --- | --- | --- | --- | --- |
| AVG. | 94.46458 | -160.51714 | 35.68131 | -0.05919 | 89.81317 | 89.91993 |
| Range | 0.00116 | 0.00089 | 0.00386 | 0.00095 | 0.00140 | 0.00270 |
| St. Dev. | 0.00044 | 0.00026 | 0.00100 | 0.00026 | 0.00037 | 0.00087 |
| RMS Tol | 0.06500 | 0.03000 | - | - | - | - |

Under Vacuum

10 Cycles (10 recordings)

TB 5-8 Locations

CSY = BENCH$CSY

| TB 5 | X | Y | Z |
| --- | --- | --- | --- |
| AVG. | 93.96599 | 102.24657 | -38.16239 |
| Range | 0.00387 | 0.00558 | 0.00067 |
| St. Dev. | 0.00136 | 0.00192 | 0.00025 |
| RMS Tol | 0.06500 | 0.03000 | - |

| TB 6 | X | Y | Z |
| --- | --- | --- | --- |
| AVG. | -94.91419 | 102.16255 | -38.22339 |
| Range | 0.00454 | 0.00545 | 0.00232 |
| St. Dev. | 0.00170 | 0.00212 | 0.00080 |
| RMS Tol | 0.06500 | 0.03000 | - |

| TB 7 | X | Y | Z |
| --- | --- | --- | --- |
| AVG. | -51.81707 | 102.43425 | 67.62683 |
| Range | 0.00385 | 0.00204 | 0.00224 |
| St. Dev. | 0.00140 | 0.00086 | 0.00083 |
| RMS Tol | 0.06500 | 0.03000 | - |

| TB 8 | X | Y | Z |
| --- | --- | --- | --- |
| AVG. | 50.76923 | 102.49155 | 67.66839 |
| Range | 0.00500 | 0.00684 | 0.00061 |
| St. Dev. | 0.00168 | 0.00221 | 0.00021 |
| RMS Tol | 0.06500 | 0.03000 | - |

Wire Location

CSY = BEAMLINE$CSY

| Stats | X | Y |
| --- | --- | --- |
| AVG. | -0.39659 | -0.18375 |
| Range | 0.01359 | 0.00060 |
| St. Dev. | 0.00500 | 0.00017 |
| Nom. | -0.40000 | -0.20000 |
| RMS Tol | 0.06500 | 0.03000 |

Transformation WIRE$CSY TO BENCH$CSY

| Stats | X | Y | Z | APA_PXPY | PA_PYPZ | PA_PXPZ |
| --- | --- | --- | --- | --- | --- | --- |
| AVG. | -0.39054 | -58.22679 | -2.12673 | -0.03399 | 89.67542 | 89.93709 |
| Range | 0.01358 | 0.00060 | 0.00103 | 0.00334 | 0.00073 | 0.00062 |
| St. Dev. | 0.00500 | 0.00017 | 0.00034 | 0.00122 | 0.00024 | 0.00020 |
| RMS Tol | 0.06500 | 0.03000 | - | - | - | - |

Locations of ¼-20 Jo-Plugs Cylinders Intersected into Plane projected 12.7mm away from surface replicating a ½” (12.7mm) offset to center of Tooling Ball.

CSY = BEAMLINE$CSY

| INTPT | X | Y | Z |
| --- | --- | --- | --- |
| 1 | -27.20874 | 227.84222 | 67.82816 |
| 2 | 26.79684 | 227.84096 | 67.86627 |
| 3 | 79.72992 | 180.18239 | -69.00653 |
| 4 | -79.92181 | 180.24115 | -69.12031 |

# Text boxes

Y+

+Z

Z+

Y+

+X

+Y

X+

Transformation