LCLS BFW WIRECARD ASSY

INSPECTION/FIDUCIAL REPORT

| Inspector: | Keith Caban |
| --- | --- |
| Responsible Engineer: | Don Schafer |
| Date: | Thursday, August 14, 2008 |
| Work Order/Charge No.: | 23800-1 |
| Serial Number: | BFW – 24 / VC – 24 / WC – 24 / BASE – 23 |
| Units of Measure: | Metric – Millimeters (mm.) |
| URL | http://www-group/met/Quality/FIDUCIAL%20REPORTS/LCLS%20BFW/BFW-24_VC-24_WC-24_BAS-23.pdf |

Coordinate System Setups

Bench CSY (BENCH$CSY)

(The only CSY in Both States of measurement - out and under vacuum)

- Plane measurement on top of Bench to set zero in Y and is the Spatial Alignment.

- Measure Pin and Diamond Pin to get centers, create a line between the two center points and this sets zero in Z and is the Planar Alignment.

- Take the symmetry point of the two pins this sets zero in X.

Bench Tooling Ball CSY (BTB$CSY)

- Create a plane using Tooling Balls 1, 2, & 3

  - This sets Zero in Y and is the spatial alignment

- Create an axis/line using TB’s 2 to 1 (+X)

  - This is the planar alignment

- TB 2 is a single point that sets zero in X, Y, Z.

BFW Tooling Ball CSY (BFWTB$CSY)

- Create a plane using Tooling Balls 5, 6, & 7

  - This sets Zero in Y and is the spatial alignment

- Create an axis/line using TB’s 6 to 5 (+X)

  - This is the planar alignment

- TB 6 is a single point that sets zero in X, Y, Z.

| View from Z+ or Downstream | View from X+ Side |
| --- | --- |

Wire CSY (WIRE$CSY)

- Back Plane (-Z) Side sets Zero in Z and is the Spatial Alignment

- Looking from +Z Side, Bottom of Card Sets Zero is the Planar Alignment

- The Intersection of the Left Side of the Card and the Bottom, sets Zero in X & Y

Wire offset from Edges per T. Nakashima’s Optical Measurement

| Wire | Wire Location | Wire Diameter |
| --- | --- | --- |
| Vertical Wire (X Location from Card) | 39.5983 | 0.0364 |
| Horizontal Wire (Y Location from Card) | 36.9967 | 0.0368 |

Support Stand CSY (BASE$CSY)

- A plane on Torque Fixture sets Zero in Y and is the Spatial Alignment

- The symmetry of the outer surfaces where “X” fine adjustments are sets zero in X.

- A line where the Torque Fixture 2X Dowel pins sets zero in Z and is the Planar Alignment.

Beamline CSY (BEAMLINE$CSY)

- Use the BENCH$CSY setup except use the Torque Fixture plate to set zero in Y and the Spatial Alignment, transform +258.67 in Y to replicate Theoretical Beamline Center

- This coordinate system will tell us where the Wire Crosshairs are to a theoretical Beam Centerline.

- The wires should be offset ~ -0.400 in X, -0.200 in Y.

Statistical Data (mm.)

Out of Vacuum

10 Cycles (10 recordings)

Card Locations  
CSY = BENCH$CSY

| STATS | X-POS CARD | X-INTPT CARD | Y-INTPT CARD |
| --- | --- | --- | --- |
| AVG. | 0.09417 | -40.08319 | -95.04640 |
| Range | 0.05010 | 0.08986 | 0.00804 |
| St. Dev. | 0.01693 | 0.02996 | 0.00230 |

Wire Location

CSY = BEAMLINE$CSY

| STATS | X WIRE | Y WIRE |
| --- | --- | --- |
| AVG. | -0.32338 | -0.21399 |
| Range | 0.04923 | 0.01160 |
| St. Dev. | 0.01655 | 0.00415 |
| RMS Tol | 0.06500 | 0.03000 |
| Nom. Loc. | -0.40000 | -0.20000 |

Transformation WIRE$CSY TO BENCH$CSY

| Stats | X | Y | Z | APA_PXPY | PA_PYPZ | PA_PXPZ |
| --- | --- | --- | --- | --- | --- | --- |
| AVG. | -0.32338 | -58.21399 | -3.01772 | -0.22983 | 90.04410 | 90.04970 |
| Range | 0.04923 | 0.01160 | 0.05577 | 0.02520 | 0.02265 | 0.01689 |
| St. Dev. | 0.01655 | 0.00415 | 0.02151 | 0.00903 | 0.00837 | 0.00599 |
| RMS Tol | 0.06500 | 0.03000 | - | - | - | - |

Transformation WIRE$CSY to BFWTB$CSY

| Stats | X | Y | Z | APA_PXPY | PA_PYPZ | PA_PXPZ |
| --- | --- | --- | --- | --- | --- | --- |
| AVG. | 93.99265 | -160.78984 | 35.71974 | -0.08492 | 89.88042 | 89.89133 |
| Range | 0.01042 | 0.00887 | 0.03191 | 0.01015 | 0.01307 | 0.01455 |
| St. Dev. | 0.00377 | 0.00303 | 0.01309 | 0.00365 | 0.00502 | 0.00492 |
| RMS Tol | 0.06500 | 0.03000 | - | - | - | - |

Under Vacuum

10 Cycles (10 recordings)

TB 5-8 Locations

CSY = BENCH$CSY

| TB 5 | X | Y | Z |
| --- | --- | --- | --- |
| AVG. | 94.66912 | 102.33031 | -38.57633 |
| Range | 0.02300 | 0.04822 | 0.01211 |
| St. Dev. | 0.00848 | 0.01775 | 0.00435 |
| RMS Tol | 0.06500 | 0.03000 | - |

| TB 6 | X | Y | Z |
| --- | --- | --- | --- |
| AVG. | -94.19133 | 102.66969 | -38.15371 |
| Range | 0.01765 | 0.02679 | 0.00344 |
| St. Dev. | 0.00598 | 0.00914 | 0.00129 |
| RMS Tol | 0.06500 | 0.03000 | - |

| TB 7 | X | Y | Z |
| --- | --- | --- | --- |
| AVG. | -50.82422 | 102.64677 | 67.57045 |
| Range | 0.01686 | 0.00587 | 0.00175 |
| St. Dev. | 0.00549 | 0.00192 | 0.00063 |
| RMS Tol | 0.06500 | 0.03000 | - |

| TB 8 | X | Y | Z |
| --- | --- | --- | --- |
| AVG. | 51.75641 | 102.47062 | 67.35305 |
| Range | 0.01399 | 0.01647 | 0.00457 |
| St. Dev. | 0.00449 | 0.00535 | 0.00147 |
| RMS Tol | 0.06500 | 0.03000 | - |

Wire Location

CSY = BEAMLINE$CSY

| STATS | X WIRE | Y WIRE |
| --- | --- | --- |
| Vac. AVG. | -0.40775 | -0.27833 |
| Range | 0.08123 | 0.01621 |
| St. Dev. | 0.02607 | 0.00510 |
| Nominal | -0.40000 | -0.20000 |
| RMS Tol | 0.06500 | 0.03000 |

Transformation WIRE$CSY TO BENCH$CSY

| Stats | X | Y | Z | APA_PXPY | PA_PYPZ | PA_PXPZ |
| --- | --- | --- | --- | --- | --- | --- |
| AVG. | -0.40447 | -58.27433 | -2.56973 | -0.19261 | 89.85507 | 89.76306 |
| Range | 0.08307 | 0.01622 | 0.02021 | 0.02260 | 0.00681 | 0.00302 |
| St. Dev. | 0.02799 | 0.00563 | 0.00634 | 0.00767 | 0.00203 | 0.00104 |
| RMS Tol | 0.06500 | 0.03000 | - | - | - | - |

Locations of ¼-20 Jo-Plugs Cylinders Intersected into Plane projected 12.7mm away from surface replicating a ½” (12.7mm) offset to center of Tooling Ball.

CSY = BEAMLINE$CSY

| INTPT | X | Y | Z |
| --- | --- | --- | --- |
| 1 | -26.98272 | 227.62923 | 68.41232 |
| 2 | 27.01037 | 227.61044 | 68.47183 |
| 3 | 80.02611 | 180.30083 | -68.49495 |
| 4 | -79.68644 | 180.31075 | -68.66787 |

# Text boxes

Y+

+Z

Z+

Y+

+X

+Y

X+

Transformation