LCLS BFW WIRECARD ASSY

INSPECTION/FIDUCIAL REPORT

| Inspector: | Keith Caban |
| --- | --- |
| Responsible Engineer: | Don Schafer |
| Date: | Tuesday, August 05, 2008 |
| Work Order/Charge No.: | 23800-1 |
| Serial Number: | BFW – 25 / VC – 25 / WC – 25 / BASE – 21 |
| Units of Measure: | Metric – Millimeters (mm.) |
| URL | http://www-group/met/Quality/FIDUCIAL%20REPORTS/LCLS%20BFW/BFW-25_VC-25_WC-25_BAS-21.pdf |

Coordinate System Setups

Bench CSY (BENCH$CSY)

(The only CSY in Both States of measurement - out and under vacuum)

- Plane measurement on top of Bench to set zero in Y and is the Spatial Alignment.

- Measure Pin and Diamond Pin to get centers, create a line between the two center points and this sets zero in Z and is the Planar Alignment.

- Take the symmetry point of the two pins this sets zero in X.

Bench Tooling Ball CSY (BTB$CSY)

- Create a plane using Tooling Balls 1, 2, & 3

  - This sets Zero in Y and is the spatial alignment

- Create an axis/line using TB’s 2 to 1 (+X)

  - This is the planar alignment

- TB 2 is a single point that sets zero in X, Y, Z.

BFW Tooling Ball CSY (BFWTB$CSY)

- Create a plane using Tooling Balls 5, 6, & 7

  - This sets Zero in Y and is the spatial alignment

- Create an axis/line using TB’s 6 to 5 (+X)

  - This is the planar alignment

- TB 6 is a single point that sets zero in X, Y, Z.

| View from Z+ or Downstream | View from X+ Side |
| --- | --- |

Wire CSY (WIRE$CSY)

- Back Plane (-Z) Side sets Zero in Z and is the Spatial Alignment

- Looking from +Z Side, Bottom of Card Sets Zero is the Planar Alignment

- The Intersection of the Left Side of the Card and the Bottom, sets Zero in X & Y

Wire offset from Edges per T. Nakashima’s Optical Measurement

| Wire | Wire Location | Wire Diameter |
| --- | --- | --- |
| Vertical Wire (X Location from Card) | 39.5940 | 0.0389 |
| Horizontal Wire (Y Location from Card) | 37.0068 | 0.0394 |

Support Stand CSY (BASE$CSY)

- A plane on Torque Fixture sets Zero in Y and is the Spatial Alignment

- The symmetry of the outer surfaces where “X” fine adjustments are sets zero in X.

- A line where the Torque Fixture 2X Dowel pins sets zero in Z and is the Planar Alignment.

Beamline CSY (BEAMLINE$CSY)

- Use the BENCH$CSY setup except use the Torque Fixture plate to set zero in Y and the Spatial Alignment, transform +258.67 in Y to replicate Theoretical Beamline Center

- This coordinate system will tell us where the Wire Crosshairs are to a theoretical Beam Centerline.

- The wires should be offset ~ -0.400 in X, -0.200 in Y.

Statistical Data (mm.)

Out of Vacuum

10 Cycles (10 recordings)

Card Locations  
CSY = BENCH$CSY

| STATS | X-POS CARD | X-INTPT CARD | Y-INTPT CARD |
| --- | --- | --- | --- |
| AVG. | 0.00566 | -40.14609 | -95.14610 |
| Range | 0.03382 | 0.04095 | 0.00394 |
| St. Dev. | 0.01058 | 0.01284 | 0.00117 |

Wire Location

CSY = BEAMLINE$CSY

| Stats | X | Y |
| --- | --- | --- |
| AVG. | -0.42275 | -0.28819 |
| Range | 0.03340 | 0.01351 |
| St. Dev. | 0.01047 | 0.00398 |
| Nom. | -0.40000 | -0.20000 |
| RMS Tol | 0.06500 | 0.03000 |

Transformation WIRE$CSY TO BENCH$CSY

| Stats | X | Y | Z | APA_PXPY | PA_PYPZ | PA_PXPZ |
| --- | --- | --- | --- | --- | --- | --- |
| AVG. | -0.42275 | -58.28819 | -2.47573 | -0.20046 | 89.79680 | 89.91017 |
| Range | 0.03340 | 0.01351 | 0.21279 | 0.01409 | 0.06267 | 0.02138 |
| St. Dev. | 0.01047 | 0.00398 | 0.06302 | 0.00430 | 0.01855 | 0.00637 |
| RMS Tol | 0.06500 | 0.03000 | - | - | - | - |

Transformation WIRE$CSY to BFWTB$CSY

| Stats | X | Y | Z | APA_PXPY | PA_PYPZ | PA_PXPZ |
| --- | --- | --- | --- | --- | --- | --- |
| AVG. | 94.49808 | -160.65437 | 35.62308 | -0.17314 | 89.89248 | 89.98426 |
| Range | 0.00158 | 0.00055 | 0.00275 | 0.00067 | 0.00199 | 0.00106 |
| St. Dev. | 0.00056 | 0.00015 | 0.00090 | 0.00021 | 0.00056 | 0.00031 |
| RMS Tol | 0.06500 | 0.03000 | - | - | - | - |

Under Vacuum

10 Cycles (10 recordings)

TB 5-8 Locations

CSY = BENCH$CSY

| TB 5 | X | Y | Z |
| --- | --- | --- | --- |
| AVG. | 94.01080 | 102.42475 | -38.32217 |
| Range | 0.01263 | 0.01663 | 0.01968 |
| St. Dev. | 0.00363 | 0.00444 | 0.00540 |
| RMS Tol | 0.06500 | 0.03000 | - |

| TB 6 | X | Y | Z |
| --- | --- | --- | --- |
| AVG. | -94.84984 | 102.57179 | -38.07493 |
| Range | 0.00907 | 0.01775 | 0.00332 |
| St. Dev. | 0.00253 | 0.00494 | 0.00103 |
| RMS Tol | 0.06500 | 0.03000 | - |

| TB 7 | X | Y | Z |
| --- | --- | --- | --- |
| AVG. | -51.59140 | 102.64753 | 67.70866 |
| Range | 0.00474 | 0.01751 | 0.00491 |
| St. Dev. | 0.00127 | 0.00477 | 0.00147 |
| RMS Tol | 0.06500 | 0.03000 | - |

| TB 8 | X | Y | Z |
| --- | --- | --- | --- |
| AVG. | 50.99040 | 102.56264 | 67.56961 |
| Range | 0.00457 | 0.01698 | 0.01783 |
| St. Dev. | 0.00125 | 0.00463 | 0.00513 |
| RMS Tol | 0.06500 | 0.03000 | - |

Wire Location

CSY = BEAMLINE$CSY

| Stats | X | Y |
| --- | --- | --- |
| AVG. | -0.42648 | -0.21768 |
| Range | 0.00671 | 0.00552 |
| St. Dev. | 0.00200 | 0.00159 |
| Nom. | -0.40000 | -0.20000 |
| RMS Tol | 0.06500 | 0.03000 |

Transformation WIRE$CSY TO BENCH$CSY

| Stats | X | Y | Z | APA_PXPY | PA_PYPZ | PA_PXPZ |
| --- | --- | --- | --- | --- | --- | --- |
| AVG. | -0.42963 | -58.11914 | -2.40861 | -0.21769 | 89.83260 | 89.90964 |
| Range | 0.00670 | 0.00558 | 0.06197 | 0.00061 | 0.01895 | 0.00656 |
| St. Dev. | 0.00200 | 0.00161 | 0.01699 | 0.00019 | 0.00513 | 0.00171 |
| RMS Tol | 0.06500 | 0.03000 | - | - | - | - |

Locations of ¼-20 Jo-Plugs Cylinders Intersected into Plane projected 12.7mm away from surface replicating a ½” (12.7mm) offset to center of Tooling Ball.

CSY = BEAMLINE$CSY

| INTPT | X | Y | Z |
| --- | --- | --- | --- |
| 1 | -26.93769 | 227.46897 | 68.60619 |
| 2 | 27.06797 | 227.46099 | 68.56112 |
| 3 | 79.80323 | 180.25730 | -68.53373 |
| 4 | -79.86524 | 180.31934 | -68.39707 |

# Text boxes

Y+

+Z

Z+

Y+

+X

+Y

X+

Transformation