LCLS BFW WIRECARD ASSY

INSPECTION/FIDUCIAL REPORT

| Inspector: | Keith Caban |
| --- | --- |
| Responsible Engineer: | Venkat Srinivasan |
| Date: | Wednesday, July 09, 2008 |
| Work Order/Charge No.: | 23800-1 |
| Serial Number: | BFW – 26 / VC – 26 / WC – 26 / BASE – 25 |
| Units of Measure: | Metric – Millimeters (mm.) |
| URL | http://www-group/met/Quality/FIDUCIAL%20REPORTS/LCLS%20BFW/BFW-26_VC-26_WC-26_BAS-25.pdf |

Coordinate System Setups

Bench CSY (BENCH$CSY)

(The only CSY in Both States of measurement - out and under vacuum)

- Plane measurement on top of Bench to set zero in Y and is the Spatial Alignment.

- Measure Pin and Diamond Pin to get centers, create a line between the two center points and this sets zero in Z and is the Planar Alignment.

- Take the symmetry point of the two pins this sets zero in X.

Bench Tooling Ball CSY (BTB$CSY)

- Create a plane using Tooling Balls 1, 2, & 3

  - This sets Zero in Y and is the spatial alignment

- Create an axis/line using TB’s 2 to 1 (+X)

  - This is the planar alignment

- TB 2 is a single point that sets zero in X, Y, Z.

BFW Tooling Ball CSY (BFWTB$CSY)

- Create a plane using Tooling Balls 5, 6, & 7

  - This sets Zero in Y and is the spatial alignment

- Create an axis/line using TB’s 6 to 5 (+X)

  - This is the planar alignment

- TB 6 is a single point that sets zero in X, Y, Z.

| View from Z+ or Downstream | View from X+ Side |
| --- | --- |

Wire CSY (WIRE$CSY)

- Back Plane (-Z) Side sets Zero in Z and is the Spatial Alignment

- Looking from +Z Side, Bottom of Card Sets Zero is the Planar Alignment

- The Intersection of the Left Side of the Card and the Bottom, sets Zero in X & Y

Wire offset from Edges per T. Nakashima’s Optical Measurement

| Wire | Wire Location | Wire Diameter |
| --- | --- | --- |
| Vertical Wire (X Location from Card) | 39.5972 | 0.0404 |
| Horizontal Wire (Y Location from Card) | 36.9945 | 0.0390 |

Support Stand CSY (BASE$CSY)

- A plane on Torque Fixture sets Zero in Y and is the Spatial Alignment

- The symmetry of the outer surfaces where “X” fine adjustments are sets zero in X.

- A line where the Torque Fixture 2X Dowel pins sets zero in Z and is the Planar Alignment.

Beamline CSY (BEAMLINE$CSY)

- Use the BENCH$CSY setup except use the Torque Fixture plate to set zero in Y and the Spatial Alignment, transform +258.67 in Y to replicate Theoretical Beamline Center

- This coordinate system will tell us where the Wire Crosshairs are to a theoretical Beam Centerline.

- The wires should be offset ~ -0.400 in X, -0.200 in Y.

Statistical Data (mm.)

Out of Vacuum

10 Cycles (10 recordings)

Card Locations  
CSY = BENCH$CSY

| STATS | X-POS CARD | X-INTPT CARD | Y-INTPT CARD |
| --- | --- | --- | --- |
| AVG. | 0.00788 | -39.99392 | -95.23768 |
| Range | 0.00189 | 0.00239 | 0.00267 |
| St. Dev. | 0.00064 | 0.00079 | 0.00090 |

Wire Location

CSY = BEAMLINE$CSY

| Stats | X | Y |
| --- | --- | --- |
| AVG. | -0.41797 | -0.22393 |
| Range | 0.00176 | 0.00290 |
| St. Dev. | 0.00063 | 0.00084 |
| Nom. | -0.40000 | -0.20000 |
| RMS Tol | 0.06500 | 0.03000 |

Transformation WIRE$CSY TO BENCH$CSY

| Stats | X | Y | Z | APA_PXPY | PA_PYPZ | PA_PXPZ |
| --- | --- | --- | --- | --- | --- | --- |
| AVG. | -0.41797 | -58.22393 | -3.33206 | 0.03816 | 90.13561 | 89.78300 |
| Range | 0.00176 | 0.00290 | 0.01125 | 0.00125 | 0.00310 | 0.00722 |
| St. Dev. | 0.00063 | 0.00084 | 0.00337 | 0.00036 | 0.00092 | 0.00215 |
| RMS Tol | 0.06500 | 0.03000 | - | - | - | - |

Transformation WIRE$CSY to BFWTB$CSY

| Stats | X | Y | Z | APA_PXPY | PA_PYPZ | PA_PXPZ |
| --- | --- | --- | --- | --- | --- | --- |
| AVG. | 94.33917 | -160.97046 | 35.88291 | 0.07000 | 89.69035 | 89.84596 |
| Range | 0.00507 | 0.00145 | 0.00392 | 0.00078 | 0.00119 | 0.00106 |
| St. Dev. | 0.00128 | 0.00040 | 0.00114 | 0.00027 | 0.00034 | 0.00033 |
| RMS Tol | 0.06500 | 0.03000 | - | - | - | - |

Under Vacuum

10 Cycles (10 recordings)

TB 5-8 Locations

CSY = BENCH$CSY

| TB 5 | X | Y | Z |
| --- | --- | --- | --- |
| AVG. | 94.26746 | 102.85791 | -37.78788 |
| Range | 0.01256 | 0.04370 | 0.01662 |
| St. Dev. | 0.00464 | 0.01300 | 0.00531 |
| RMS Tol | 0.06500 | 0.03000 | - |

| TB 6 | X | Y | Z |
| --- | --- | --- | --- |
| AVG. | -94.58681 | 103.11477 | -37.65227 |
| Range | 0.01390 | 0.03740 | 0.02518 |
| St. Dev. | 0.00473 | 0.01263 | 0.00909 |
| RMS Tol | 0.06500 | 0.03000 | - |

| TB 7 | X | Y | Z |
| --- | --- | --- | --- |
| AVG. | -51.40488 | 102.69157 | 68.14024 |
| Range | 0.01121 | 0.02754 | 0.01526 |
| St. Dev. | 0.00383 | 0.01039 | 0.00557 |
| RMS Tol | 0.06500 | 0.03000 | - |

| TB 8 | X | Y | Z |
| --- | --- | --- | --- |
| AVG. | 51.18548 | 102.56069 | 68.07580 |
| Range | 0.01126 | 0.03440 | 0.01462 |
| St. Dev. | 0.00385 | 0.01069 | 0.00415 |
| RMS Tol | 0.06500 | 0.03000 | - |

Wire Location

CSY = BEAMLINE$CSY

| Stats | X | Y |
| --- | --- | --- |
| AVG. | -0.42936 | -0.20390 |
| Range | 0.03842 | 0.02429 |
| St. Dev. | 0.01176 | 0.00906 |
| Nom. | -0.40000 | -0.20000 |
| RMS Tol | 0.06500 | 0.03000 |

Transformation WIRE$CSY TO BENCH$CSY

| Stats | X | Y | Z | APA_PXPY | PA_PYPZ | PA_PXPZ |
| --- | --- | --- | --- | --- | --- | --- |
| AVG. | -0.43856 | -58.10731 | -2.39128 | -0.00718 | 89.88801 | 89.80496 |
| Range | 0.03842 | 0.02428 | 0.00701 | 0.01529 | 0.00069 | 0.01265 |
| St. Dev. | 0.01294 | 0.00944 | 0.00222 | 0.00525 | 0.00018 | 0.00434 |
| RMS Tol | 0.06500 | 0.03000 | - | - | - | - |

Locations of ¼-20 Jo-Plugs Cylinders Intersected into Plane projected 12.7mm away from surface replicating a ½” (12.7mm) offset to center of Tooling Ball.

CSY = BEAMLINE$CSY

| INTPT | X | Y | Z |
| --- | --- | --- | --- |
| 1 | -26.72052 | 227.63543 | 68.31481 |
| 2 | 27.23137 | 227.65552 | 68.23792 |
| 3 | 79.88810 | 180.03909 | -68.79429 |
| 4 | -79.79443 | 180.05734 | -68.56294 |

# Text boxes

Y+

+Z

Z+

Y+

+X

+Y

X+

Transformation