LCLS BFW WIRECARD ASSY

INSPECTION/FIDUCIAL REPORT

| Inspector: | Keith Caban |
| --- | --- |
| Responsible Engineer: | Don Schafer |
| Date: | Thursday, July 31, 2008 |
| Work Order/Charge No.: | 23800-1 |
| Serial Number: | BFW – 27 / VC – 27 / WC – 27 / BASE – 02 |
| Units of Measure: | Metric – Millimeters (mm.) |
| URL | http://www-group/met/Quality/FIDUCIAL%20REPORTS/LCLS%20BFW/BFW-27_VC-27_WC-27_BAS-02.pdf |

Coordinate System Setups

Bench CSY (BENCH$CSY)

(The only CSY in Both States of measurement - out and under vacuum)

- Plane measurement on top of Bench to set zero in Y and is the Spatial Alignment.

- Measure Pin and Diamond Pin to get centers, create a line between the two center points and this sets zero in Z and is the Planar Alignment.

- Take the symmetry point of the two pins this sets zero in X.

Bench Tooling Ball CSY (BTB$CSY)

- Create a plane using Tooling Balls 1, 2, & 3

  - This sets Zero in Y and is the spatial alignment

- Create an axis/line using TB’s 2 to 1 (+X)

  - This is the planar alignment

- TB 2 is a single point that sets zero in X, Y, Z.

BFW Tooling Ball CSY (BFWTB$CSY)

- Create a plane using Tooling Balls 5, 6, & 7

  - This sets Zero in Y and is the spatial alignment

- Create an axis/line using TB’s 6 to 5 (+X)

  - This is the planar alignment

- TB 6 is a single point that sets zero in X, Y, Z.

| View from Z+ or Downstream | View from X+ Side |
| --- | --- |

Wire CSY (WIRE$CSY)

- Back Plane (-Z) Side sets Zero in Z and is the Spatial Alignment

- Looking from +Z Side, Bottom of Card Sets Zero is the Planar Alignment

- The Intersection of the Left Side of the Card and the Bottom, sets Zero in X & Y

Wire offset from Edges per T. Nakashima’s Optical Measurement

| Wire | Wire Location | Wire Diameter |
| --- | --- | --- |
| Vertical Wire (X Location from Card) | 39.5937 | 0.0384 |
| Horizontal Wire (Y Location from Card) | 37.0025 | 0.0394 |

Support Stand CSY (BASE$CSY)

- A plane on Torque Fixture sets Zero in Y and is the Spatial Alignment

- The symmetry of the outer surfaces where “X” fine adjustments are sets zero in X.

- A line where the Torque Fixture 2X Dowel pins sets zero in Z and is the Planar Alignment.

Beamline CSY (BEAMLINE$CSY)

- Use the BENCH$CSY setup except use the Torque Fixture plate to set zero in Y and the Spatial Alignment, transform +258.67 in Y to replicate Theoretical Beamline Center

- This coordinate system will tell us where the Wire Crosshairs are to a theoretical Beam Centerline.

- The wires should be offset ~ -0.400 in X, -0.200 in Y.

Statistical Data (mm.)

Out of Vacuum

10 Cycles (10 recordings)

Card Locations  
CSY = BENCH$CSY

| STATS | X-POS CARD | X-INTPT CARD | Y-INTPT CARD |
| --- | --- | --- | --- |
| AVG. | -0.05050 | -40.22211 | -95.09175 |
| Range | 0.05699 | 0.09291 | 0.00339 |
| St. Dev. | 0.02122 | 0.03322 | 0.00108 |

Wire Location

CSY = BEAMLINE$CSY

| Stats | X | Y |
| --- | --- | --- |
| AVG. | -0.48460 | -0.23782 |
| Range | 0.05701 | 0.00888 |
| St. Dev. | 0.02130 | 0.00312 |
| Nom. | -0.40000 | -0.20000 |
| RMS Tol | 0.06500 | 0.03000 |

Transformation WIRE$CSY TO BENCH$CSY

| Stats | X | Y | Z | APA_PXPY | PA_PYPZ | PA_PXPZ |
| --- | --- | --- | --- | --- | --- | --- |
| AVG. | -0.48460 | -58.23782 | -2.88739 | -0.20666 | 89.85741 | 89.87540 |
| Range | 0.05701 | 0.00888 | 0.03408 | 0.02062 | 0.00652 | 0.00392 |
| St. Dev. | 0.02130 | 0.00312 | 0.01215 | 0.00731 | 0.00224 | 0.00112 |
| RMS Tol | 0.06500 | 0.03000 | - | - | - | - |

Transformation WIRE$CSY to BFWTB$CSY

| Stats | X | Y | Z | APA_PXPY | PA_PYPZ | PA_PXPZ |
| --- | --- | --- | --- | --- | --- | --- |
| AVG. | 94.35796 | -160.80303 | 35.80487 | 0.02849 | 89.76585 | 89.88179 |
| Range | 0.05102 | 0.00748 | 0.01173 | 0.01247 | 0.00228 | 0.00216 |
| St. Dev. | 0.01572 | 0.00223 | 0.00366 | 0.00335 | 0.00064 | 0.00068 |
| RMS Tol | 0.06500 | 0.03000 | - | - | - | - |

Under Vacuum

10 Cycles (10 recordings)

TB 5-8 Locations

CSY = BENCH$CSY

| TB 5 | X | Y | Z |
| --- | --- | --- | --- |
| AVG. | 94.78031 | 102.30443 | -38.05835 |
| Range | 0.03475 | 0.07589 | 0.01918 |
| St. Dev. | 0.01091 | 0.02354 | 0.00560 |
| RMS Tol | 0.06500 | 0.03000 | - |

| TB 6 | X | Y | Z |
| --- | --- | --- | --- |
| AVG. | -94.10272 | 103.02251 | -38.51155 |
| Range | 0.03422 | 0.04123 | 0.01397 |
| St. Dev. | 0.01074 | 0.01295 | 0.00446 |
| RMS Tol | 0.06500 | 0.03000 | - |

| TB 7 | X | Y | Z |
| --- | --- | --- | --- |
| AVG. | -51.22147 | 102.81982 | 67.43276 |
| Range | 0.03746 | 0.00342 | 0.01443 |
| St. Dev. | 0.01144 | 0.00097 | 0.00457 |
| RMS Tol | 0.06500 | 0.03000 | - |

| TB 8 | X | Y | Z |
| --- | --- | --- | --- |
| AVG. | 51.35127 | 102.42321 | 67.66543 |
| Range | 0.03756 | 0.06427 | 0.01661 |
| St. Dev. | 0.01152 | 0.01967 | 0.00518 |
| RMS Tol | 0.06500 | 0.03000 | - |

Wire Location

CSY = BEAMLINE$CSY

| Stats | X | Y |
| --- | --- | --- |
| AVG. | -0.42826 | -0.16652 |
| Range | 0.13520 | 0.02194 |
| St. Dev. | 0.04207 | 0.00668 |
| Nom. | -0.40000 | -0.20000 |
| RMS Tol | 0.06500 | 0.03000 |

Transformation WIRE$CSY TO BENCH$CSY

| Stats | X | Y | Z | APA_PXPY | PA_PYPZ | PA_PXPZ |
| --- | --- | --- | --- | --- | --- | --- |
| AVG. | -0.43955 | -58.15035 | -2.53825 | -0.19018 | 89.78700 | 90.01932 |
| Range | 0.13519 | 0.02190 | 0.00656 | 0.03551 | 0.00806 | 0.00163 |
| St. Dev. | 0.04207 | 0.00667 | 0.00201 | 0.01106 | 0.00244 | 0.00045 |
| RMS Tol | 0.06500 | 0.03000 | - | - | - | - |

Locations of ¼-20 Jo-Plugs Cylinders Intersected into Plane projected 12.7mm away from surface replicating a ½” (12.7mm) offset to center of Tooling Ball.

CSY = BEAMLINE$CSY

| INTPT | X | Y | Z |
| --- | --- | --- | --- |
| 1 | -27.07100 | 227.68239 | 68.16338 |
| 2 | 26.93968 | 227.69649 | 68.14872 |
| 3 | 79.75530 | 180.19309 | -68.81245 |
| 4 | -79.97013 | 180.15819 | -68.75423 |

# Text boxes

Y+

+Z

Z+

Y+

+X

+Y

X+

Transformation