LCLS BFW WIRECARD ASSY

INSPECTION/FIDUCIAL REPORT

| Inspector: | Keith Caban |
| --- | --- |
| Responsible Engineer: | Venkat Srinivasan |
| Date: | Friday, July 18, 2008 |
| Work Order/Charge No.: | 23800-1 |
| Serial Number: | BFW – 30 / VC – 30 / WC – 30 / BASE – 22 |
| Units of Measure: | Metric – Millimeters (mm.) |
| URL | http://www-group/met/Quality/FIDUCIAL%20REPORTS/LCLS%20BFW/BFW-30_VC-30_WC-30_BAS-22.pdf |

Coordinate System Setups

Bench CSY (BENCH$CSY)

(The only CSY in Both States of measurement - out and under vacuum)

- Plane measurement on top of Bench to set zero in Y and is the Spatial Alignment.

- Measure Pin and Diamond Pin to get centers, create a line between the two center points and this sets zero in Z and is the Planar Alignment.

- Take the symmetry point of the two pins this sets zero in X.

Bench Tooling Ball CSY (BTB$CSY)

- Create a plane using Tooling Balls 1, 2, & 3

  - This sets Zero in Y and is the spatial alignment

- Create an axis/line using TB’s 2 to 1 (+X)

  - This is the planar alignment

- TB 2 is a single point that sets zero in X, Y, Z.

BFW Tooling Ball CSY (BFWTB$CSY)

- Create a plane using Tooling Balls 5, 6, & 7

  - This sets Zero in Y and is the spatial alignment

- Create an axis/line using TB’s 6 to 5 (+X)

  - This is the planar alignment

- TB 6 is a single point that sets zero in X, Y, Z.

| View from Z+ or Downstream | View from X+ Side |
| --- | --- |

Wire CSY (WIRE$CSY)

- Back Plane (-Z) Side sets Zero in Z and is the Spatial Alignment

- Looking from +Z Side, Bottom of Card Sets Zero is the Planar Alignment

- The Intersection of the Left Side of the Card and the Bottom, sets Zero in X & Y

Wire offset from Edges per T. Nakashima’s Optical Measurement

| Wire | Wire Location | Wire Diameter |
| --- | --- | --- |
| Vertical Wire (X Location from Card) | 39.6085 | 0.0358 |
| Horizontal Wire (Y Location from Card) | 37.0157 | 0.0375 |

Support Stand CSY (BASE$CSY)

- A plane on Torque Fixture sets Zero in Y and is the Spatial Alignment

- The symmetry of the outer surfaces where “X” fine adjustments are sets zero in X.

- A line where the Torque Fixture 2X Dowel pins sets zero in Z and is the Planar Alignment.

Beamline CSY (BEAMLINE$CSY)

- Use the BENCH$CSY setup except use the Torque Fixture plate to set zero in Y and the Spatial Alignment, transform +258.67 in Y to replicate Theoretical Beamline Center

- This coordinate system will tell us where the Wire Crosshairs are to a theoretical Beam Centerline.

- The wires should be offset ~ -0.400 in X, -0.200 in Y.

Statistical Data (mm.)

Out of Vacuum

10 Cycles (10 recordings)

Card Locations  
CSY = BENCH$CSY

| STATS | X-POS CARD | X-INTPT CARD | Y-INTPT CARD |
| --- | --- | --- | --- |
| 0.04187 | -40.06565 | -95.11445 | 0.04187 |
| 0.00150 | 0.00173 | 0.00149 | 0.00150 |
| 0.00056 | 0.00069 | 0.00045 | 0.00056 |

Wire Location

CSY = BEAMLINE$CSY

| Stats | X | Y |
| --- | --- | --- |
| AVG. | -0.37937 | -0.18985 |
| Range | 0.00154 | 0.00167 |
| St. Dev. | 0.00058 | 0.00045 |
| Nom. | -0.40000 | -0.20000 |
| RMS Tol | 0.06500 | 0.03000 |

Transformation WIRE$CSY TO BENCH$CSY

| Stats | X | Y | Z | APA_PXPY | PA_PYPZ | PA_PXPZ |
| --- | --- | --- | --- | --- | --- | --- |
| AVG. | -0.37937 | -58.18985 | -2.81786 | -0.11533 | 90.06411 | 90.00551 |
| Range | 0.00154 | 0.00167 | 0.00691 | 0.00082 | 0.00258 | 0.00342 |
| St. Dev. | 0.00058 | 0.00045 | 0.00253 | 0.00027 | 0.00110 | 0.00114 |
| RMS Tol | 0.06500 | 0.03000 | - | - | - | - |

Transformation WIRE$CSY to BFWTB$CSY

| Stats | X | Y | Z | APA_PXPY | PA_PYPZ | PA_PXPZ |
| --- | --- | --- | --- | --- | --- | --- |
| AVG. | 94.53492 | -160.97775 | 35.62955 | -0.07111 | 89.89784 | 89.89722 |
| Range | 0.00310 | 0.00100 | 0.00246 | 0.00097 | 0.00214 | 0.00344 |
| St. Dev. | 0.00097 | 0.00031 | 0.00081 | 0.00025 | 0.00076 | 0.00101 |
| RMS Tol | 0.06500 | 0.03000 | - | - | - | - |

Under Vacuum

10 Cycles (10 recordings)

TB 5-8 Locations

CSY = BENCH$CSY

| TB 5 | X | Y | Z |
| --- | --- | --- | --- |
| AVG. | 94.10355 | 102.78085 | -37.86279 |
| Range | 0.00175 | 0.00498 | 0.00114 |
| St. Dev. | 0.00059 | 0.00172 | 0.00031 |
| RMS Tol | 0.06500 | 0.03000 | - |

| TB 6 | X | Y | Z |
| --- | --- | --- | --- |
| AVG. | -94.77206 | 102.99674 | -38.10528 |
| Range | 0.00199 | 0.00139 | 0.00203 |
| St. Dev. | 0.00063 | 0.00041 | 0.00059 |
| RMS Tol | 0.06500 | 0.03000 | - |

| TB 7 | X | Y | Z |
| --- | --- | --- | --- |
| AVG. | -51.79541 | 102.73020 | 67.78382 |
| Range | 0.00138 | 0.00394 | 0.00202 |
| St. Dev. | 0.00040 | 0.00113 | 0.00062 |
| RMS Tol | 0.06500 | 0.03000 | - |

| TB 8 | X | Y | Z |
| --- | --- | --- | --- |
| AVG. | 50.78192 | 102.63268 | 67.91857 |
| Range | 0.00145 | 0.00620 | 0.00070 |
| St. Dev. | 0.00046 | 0.00198 | 0.00021 |
| RMS Tol | 0.06500 | 0.03000 | - |

Wire Location

CSY = BEAMLINE$CSY

| Stats | X | Y |
| --- | --- | --- |
| AVG. | -0.46691 | -0.21456 |
| Range | 0.00299 | 0.00340 |
| St. Dev. | 0.00112 | 0.00104 |
| Nom. | -0.40000 | -0.20000 |
| RMS Tol | 0.06500 | 0.03000 |

Transformation WIRE$CSY TO BENCH$CSY

| Stats | X | Y | Z | APA_PXPY | PA_PYPZ | PA_PXPZ |
| --- | --- | --- | --- | --- | --- | --- |
| AVG. | -0.46682 | -58.16152 | -2.68581 | -0.13651 | 90.01645 | 89.97089 |
| Range | 0.00298 | 0.00340 | 0.00505 | 0.00157 | 0.00164 | 0.00086 |
| St. Dev. | 0.00112 | 0.00104 | 0.00153 | 0.00057 | 0.00050 | 0.00026 |
| RMS Tol | 0.06500 | 0.03000 | - | - | - | - |

Locations of ¼-20 Jo-Plugs Cylinders Intersected into Plane projected 12.7mm away from surface replicating a ½” (12.7mm) offset to center of Tooling Ball.

CSY = BEAMLINE$CSY

| INTPT | X | Y | Z |
| --- | --- | --- | --- |
| 1 | -27.02011 | 227.55721 | 68.43811 |
| 2 | 26.97205 | 227.54475 | 68.41101 |
| 3 | 79.81137 | 180.28198 | -68.62541 |
| 4 | -79.88728 | 180.18823 | -68.55206 |

# Text boxes

Y+

+Z

Z+

Y+

+X

+Y

X+

Transformation