LCLS BFW WIRECARD ASSY

INSPECTION/FIDUCIAL REPORT

| Inspector: | Keith Caban |
| --- | --- |
| Responsible Engineer: | Venkat Srinivasan |
| Date: | Wednesday, July 02, 2008 |
| Work Order/Charge No.: | 23800-1 |
| Serial Number: | BFW – 32 / VC – 32 / WC – 32 / BASE – 12 |
| Units of Measure: | Metric – Millimeters (mm.) |
| URL | http://www-group/met/Quality/FIDUCIAL%20REPORTS/LCLS%20BFW/BFW-32_VC-32_WC-32_BAS-12.pdf |

Coordinate System Setups

Bench CSY (BENCH$CSY)

(The only CSY in Both States of measurement - out and under vacuum)

- Plane measurement on top of Bench to set zero in Y and is the Spatial Alignment.

- Measure Pin and Diamond Pin to get centers, create a line between the two center points and this sets zero in Z and is the Planar Alignment.

- Take the symmetry point of the two pins this sets zero in X.

Bench Tooling Ball CSY (BTB$CSY)

- Create a plane using Tooling Balls 1, 2, & 3

  - This sets Zero in Y and is the spatial alignment

- Create an axis/line using TB’s 2 to 1 (+X)

  - This is the planar alignment

- TB 2 is a single point that sets zero in X, Y, Z.

BFW Tooling Ball CSY (BFWTB$CSY)

- Create a plane using Tooling Balls 5, 6, & 7

  - This sets Zero in Y and is the spatial alignment

- Create an axis/line using TB’s 6 to 5 (+X)

  - This is the planar alignment

- TB 6 is a single point that sets zero in X, Y, Z.

| View from Z+ or Downstream | View from X+ Side |
| --- | --- |

Wire CSY (WIRE$CSY)

- Back Plane (-Z) Side sets Zero in Z and is the Spatial Alignment

- Looking from +Z Side, Bottom of Card Sets Zero is the Planar Alignment

- The Intersection of the Left Side of the Card and the Bottom, sets Zero in X & Y

Wire offset from Edges per T. Nakashima’s Optical Measurement

| Wire | Wire Location | Wire Diameter |
| --- | --- | --- |
| Vertical Wire (X Location from Card) | 39.6181 | 0.0392 |
| Horizontal Wire (Y Location from Card) | 37.0123 | 0.0374 |

Support Stand CSY (BASE$CSY)

- A plane on Torque Fixture sets Zero in Y and is the Spatial Alignment

- The symmetry of the outer surfaces where “X” fine adjustments are sets zero in X.

- A line where the Torque Fixture 2X Dowel pins sets zero in Z and is the Planar Alignment.

Beamline CSY (BEAMLINE$CSY)

- Use the BENCH$CSY setup except use the Torque Fixture plate to set zero in Y and the Spatial Alignment, transform +258.67 in Y to replicate Theoretical Beamline Center

- This coordinate system will tell us where the Wire Crosshairs are to a theoretical Beam Centerline.

- The wires should be offset ~ -0.400 in X, -0.200 in Y.

Statistical Data (mm.)

Out of Vacuum

10 Cycles (10 recordings)

Card Locations  
CSY = BENCH$CSY

| STATS | X-POS CARD | X-INTPT CARD | Y-INTPT CARD |
| --- | --- | --- | --- |
| AVG. | 0.00295 | -40.19540 | -95.08207 |
| Range | 0.01822 | 0.03186 | 0.02421 |
| St. Dev. | 0.00567 | 0.00995 | 0.00717 |

Wire Location

CSY = BEAMLINE$CSY

| Stats | X | Y |
| --- | --- | --- |
| AVG. | -0.42134 | -0.24463 |
| Range | 0.01793 | 0.01400 |
| St. Dev. | 0.00563 | 0.00417 |
| Nom. | -0.40000 | -0.20000 |
| RMS Tol | 0.06500 | 0.03000 |

Transformation WIRE$CSY TO BENCH$CSY

| Stats | X | Y | Z | APA_PXPY | PA_PYPZ | PA_PXPZ |
| --- | --- | --- | --- | --- | --- | --- |
| AVG. | -0.42134 | -58.24463 | -3.18970 | -0.24234 | 90.07620 | 89.83903 |
| Range | 0.01793 | 0.01400 | 0.01496 | 0.02334 | 0.02126 | 0.01308 |
| St. Dev. | 0.00563 | 0.00417 | 0.00542 | 0.00699 | 0.00826 | 0.00469 |
| RMS Tol | 0.06500 | 0.03000 | - | - | - | - |

Transformation WIRE$CSY to BFWTB$CSY

| Stats | X | Y | Z | APA_PXPY | PA_PYPZ | PA_PXPZ |
| --- | --- | --- | --- | --- | --- | --- |
| AVG. | 94.38405 | -161.37700 | 35.81980 | 0.01934 | 89.78127 | 89.95425 |
| Range | 0.00977 | 0.01420 | 0.01124 | 0.02132 | 0.02066 | 0.01140 |
| St. Dev. | 0.00290 | 0.00417 | 0.00391 | 0.00632 | 0.00787 | 0.00431 |
| RMS Tol | 0.06500 | 0.03000 | - | - | - | - |

Under Vacuum

10 Cycles (10 recordings)

TB 5-8 Locations

CSY = BENCH$CSY

| TB 5 | X | Y | Z |
| --- | --- | --- | --- |
| AVG. | 94.26221 | 103.07479 | -38.32306 |
| Range | 0.01680 | 0.08378 | 0.01269 |
| St. Dev. | 0.00610 | 0.02945 | 0.00465 |
| RMS Tol | 0.06500 | 0.03000 | - |

| TB 6 | X | Y | Z |
| --- | --- | --- | --- |
| AVG. | -94.62094 | 103.44402 | -37.97189 |
| Range | 0.01689 | 0.04630 | 0.00751 |
| St. Dev. | 0.00614 | 0.01668 | 0.00269 |
| RMS Tol | 0.06500 | 0.03000 | - |

| TB 7 | X | Y | Z |
| --- | --- | --- | --- |
| AVG. | -51.28821 | 103.17531 | 67.79101 |
| Range | 0.01429 | 0.02427 | 0.00870 |
| St. Dev. | 0.00516 | 0.00860 | 0.00310 |
| RMS Tol | 0.06500 | 0.03000 | - |

| TB 8 | X | Y | Z |
| --- | --- | --- | --- |
| AVG. | 51.29324 | 102.97242 | 67.58966 |
| Range | 0.01471 | 0.04682 | 0.01110 |
| St. Dev. | 0.00527 | 0.01656 | 0.00404 |
| RMS Tol | 0.06500 | 0.03000 | - |

Wire Location

CSY = BEAMLINE$CSY

| Stats | X | Y |
| --- | --- | --- |
| AVG. | -0.42055 | -0.21848 |
| Range | 0.12720 | 0.01594 |
| St. Dev. | 0.04381 | 0.00542 |
| Nom. | -0.40000 | -0.20000 |
| RMS Tol | 0.06500 | 0.03000 |

Transformation WIRE$CSY TO BENCH$CSY

| Stats | X | Y | Z | APA_PXPY | PA_PYPZ | PA_PXPZ |
| --- | --- | --- | --- | --- | --- | --- |
| AVG. | -0.42241 | -58.18286 | -2.48824 | -0.15157 | 89.79617 | 89.88313 |
| Range | 0.12720 | 0.01594 | 0.00493 | 0.03946 | 0.00414 | 0.00172 |
| St. Dev. | 0.04521 | 0.00551 | 0.00137 | 0.01399 | 0.00138 | 0.00066 |
| RMS Tol | 0.06500 | 0.03000 | - | - | - | - |

Locations of ¼-20 Jo-Plugs Cylinders Intersected into Plane projected 12.7mm away from surface replicating a ½” (12.7mm) offset to center of Tooling Ball.

CSY = BEAMLINE$CSY

| INTPT | X | Y | Z |
| --- | --- | --- | --- |
| 1 | -26.93387 | 227.53784 | 68.64131 |
| 2 | 27.10308 | 227.53542 | 68.61697 |
| 3 | 79.86901 | 180.30717 | -68.48035 |
| 4 | -79.77753 | 180.32195 | -68.40561 |

# Text boxes

Y+

+Z

Z+

Y+

+X

+Y

X+

Transformation