LCLS BFW WIRECARD ASSY

INSPECTION/FIDUCIAL REPORT

| Inspector: | Keith Caban |
| --- | --- |
| Responsible Engineer: | Don Schafer |
| Date: | Monday, August 04, 2008 |
| Work Order/Charge No.: | 23800-1 |
| Serial Number: | BFW – 33 / VC – 33 / WC – 33 / BASE – 34 |
| Units of Measure: | Metric – Millimeters (mm.) |
| URL | http://www-group/met/Quality/FIDUCIAL%20REPORTS/LCLS%20BFW/BFW-33_VC-33_WC-33_BAS-34.pdf |

Coordinate System Setups

Bench CSY (BENCH$CSY)

(The only CSY in Both States of measurement - out and under vacuum)

- Plane measurement on top of Bench to set zero in Y and is the Spatial Alignment.

- Measure Pin and Diamond Pin to get centers, create a line between the two center points and this sets zero in Z and is the Planar Alignment.

- Take the symmetry point of the two pins this sets zero in X.

Bench Tooling Ball CSY (BTB$CSY)

- Create a plane using Tooling Balls 1, 2, & 3

  - This sets Zero in Y and is the spatial alignment

- Create an axis/line using TB’s 2 to 1 (+X)

  - This is the planar alignment

- TB 2 is a single point that sets zero in X, Y, Z.

BFW Tooling Ball CSY (BFWTB$CSY)

- Create a plane using Tooling Balls 5, 6, & 7

  - This sets Zero in Y and is the spatial alignment

- Create an axis/line using TB’s 6 to 5 (+X)

  - This is the planar alignment

- TB 6 is a single point that sets zero in X, Y, Z.

| View from Z+ or Downstream | View from X+ Side |
| --- | --- |

Wire CSY (WIRE$CSY)

- Back Plane (-Z) Side sets Zero in Z and is the Spatial Alignment

- Looking from +Z Side, Bottom of Card Sets Zero is the Planar Alignment

- The Intersection of the Left Side of the Card and the Bottom, sets Zero in X & Y

Wire offset from Edges per T. Nakashima’s Optical Measurement

| Wire | Wire Location | Wire Diameter |
| --- | --- | --- |
| Vertical Wire (X Location from Card) | 39.6018 | 0.0369 |
| Horizontal Wire (Y Location from Card) | 37.0055 | 0.0368 |

Support Stand CSY (BASE$CSY)

- A plane on Torque Fixture sets Zero in Y and is the Spatial Alignment

- The symmetry of the outer surfaces where “X” fine adjustments are sets zero in X.

- A line where the Torque Fixture 2X Dowel pins sets zero in Z and is the Planar Alignment.

Beamline CSY (BEAMLINE$CSY)

- Use the BENCH$CSY setup except use the Torque Fixture plate to set zero in Y and the Spatial Alignment, transform +258.67 in Y to replicate Theoretical Beamline Center

- This coordinate system will tell us where the Wire Crosshairs are to a theoretical Beam Centerline.

- The wires should be offset ~ -0.400 in X, -0.200 in Y.

Statistical Data (mm.)

Out of Vacuum

10 Cycles (10 recordings)

Card Locations  
CSY = BENCH$CSY

| STATS | X-POS CARD | X-INTPT CARD | Y-INTPT CARD |
| --- | --- | --- | --- |
| AVG. | 0.03437 | -40.13921 | -95.02669 |
| Range | 0.02185 | 0.03837 | 0.00486 |
| St. Dev. | 0.00806 | 0.01226 | 0.00166 |

Wire Location

CSY = BEAMLINE$CSY

| Stats | X | Y |
| --- | --- | --- |
| AVG. | -0.39323 | -0.17868 |
| Range | 0.02191 | 0.00333 |
| St. Dev. | 0.00807 | 0.00108 |
| Nom. | -0.40000 | -0.20000 |
| RMS Tol | 0.06500 | 0.03000 |

Transformation WIRE$CSY TO BENCH$CSY

| Stats | X | Y | Z | APA_PXPY | PA_PYPZ | PA_PXPZ |
| --- | --- | --- | --- | --- | --- | --- |
| AVG. | -0.39323 | -58.17868 | -3.34756 | -0.21366 | 90.18786 | 90.01592 |
| Range | 0.02191 | 0.00333 | 0.00733 | 0.01265 | 0.00104 | 0.00330 |
| St. Dev. | 0.00807 | 0.00108 | 0.00269 | 0.00405 | 0.00033 | 0.00107 |
| RMS Tol | 0.06500 | 0.03000 | - | - | - | - |

Transformation WIRE$CSY to BFWTB$CSY

| Stats | X | Y | Z | APA_PXPY | PA_PYPZ | PA_PXPZ |
| --- | --- | --- | --- | --- | --- | --- |
| AVG. | 94.33475 | -160.78225 | 35.66280 | -0.04877 | 89.87981 | 90.02615 |
| Range | 0.00093 | 0.00057 | 0.00480 | 0.00766 | 0.00151 | 0.00095 |
| St. Dev. | 0.00028 | 0.00017 | 0.00152 | 0.00218 | 0.00042 | 0.00030 |
| RMS Tol | 0.06500 | 0.03000 | - | - | - | - |

Under Vacuum

10 Cycles (10 recordings)

TB 5-8 Locations

CSY = BENCH$CSY

| TB 5 | X | Y | Z |
| --- | --- | --- | --- |
| AVG. | 94.51359 | 102.44815 | -38.01897 |
| Range | 0.01106 | 0.02515 | 0.01779 |
| St. Dev. | 0.00332 | 0.00768 | 0.00607 |
| RMS Tol | 0.06500 | 0.03000 | - |

| TB 6 | X | Y | Z |
| --- | --- | --- | --- |
| AVG. | -94.35886 | 103.06203 | -37.73394 |
| Range | 0.01082 | 0.03649 | 0.01188 |
| St. Dev. | 0.00319 | 0.01079 | 0.00488 |
| RMS Tol | 0.06500 | 0.03000 | - |

| TB 7 | X | Y | Z |
| --- | --- | --- | --- |
| AVG. | -51.06389 | 102.77478 | 68.03498 |
| Range | 0.02244 | 0.01930 | 0.00687 |
| St. Dev. | 0.00639 | 0.00701 | 0.00264 |
| RMS Tol | 0.06500 | 0.03000 | - |

| TB 8 | X | Y | Z |
| --- | --- | --- | --- |
| AVG. | 51.50677 | 102.44445 | 67.87541 |
| Range | 0.02223 | 0.01541 | 0.01105 |
| St. Dev. | 0.00632 | 0.00496 | 0.00366 |
| RMS Tol | 0.06500 | 0.03000 | - |

Wire Location

CSY = BEAMLINE$CSY

| Stats | X | Y |
| --- | --- | --- |
| AVG. | -0.49417 | -0.16789 |
| Range | 0.06444 | 0.01087 |
| St. Dev. | 0.01792 | 0.00372 |
| Nom. | -0.40000 | -0.20000 |
| RMS Tol | 0.06500 | 0.03000 |

Transformation WIRE$CSY TO BENCH$CSY

| Stats | X | Y | Z | APA_PXPY | PA_PYPZ | PA_PXPZ |
| --- | --- | --- | --- | --- | --- | --- |
| AVG. | -0.49384 | -58.07556 | -2.43546 | -0.23606 | 89.95975 | 89.93908 |
| Range | 0.06443 | 0.01086 | 0.00680 | 0.01869 | 0.00255 | 0.00858 |
| St. Dev. | 0.01791 | 0.00372 | 0.00251 | 0.00523 | 0.00089 | 0.00327 |
| RMS Tol | 0.06500 | 0.03000 | - | - | - | - |

Locations of ¼-20 Jo-Plugs Cylinders Intersected into Plane projected 12.7mm away from surface replicating a ½” (12.7mm) offset to center of Tooling Ball.

CSY = BEAMLINE$CSY

| INTPT | X | Y | Z |
| --- | --- | --- | --- |
| 1 | -27.09467 | 227.48872 | 68.56493 |
| 2 | 26.93689 | 227.48698 | 68.53056 |
| 3 | 79.71340 | 180.29823 | -68.54884 |
| 4 | -79.97711 | 180.23710 | -68.43563 |

# Text boxes

Y+

+Z

Z+

Y+

+X

+Y

X+

Transformation