LCLS BFW WIRECARD ASSY

INSPECTION/FIDUCIAL REPORT

| Inspector: | Keith Caban |
| --- | --- |
| Responsible Engineer: | Don Schafer |
| Date: | Tuesday, August 19, 2008 |
| Work Order/Charge No.: | 23800-1 |
| Serial Number: | BFW – 36 / VC – 36 / WC – 36 / BASE – 36 |
| Units of Measure: | Metric – Millimeters (mm.) |
| URL | http://www-group/met/Quality/FIDUCIAL%20REPORTS/LCLS%20BFW/BFW-36_VC-36_WC-36_BAS-36.pdf |

Coordinate System Setups

Bench CSY (BENCH$CSY)

(The only CSY in Both States of measurement - out and under vacuum)

- Plane measurement on top of Bench to set zero in Y and is the Spatial Alignment.

- Measure Pin and Diamond Pin to get centers, create a line between the two center points and this sets zero in Z and is the Planar Alignment.

- Take the symmetry point of the two pins this sets zero in X.

Bench Tooling Ball CSY (BTB$CSY)

- Create a plane using Tooling Balls 1, 2, & 3

  - This sets Zero in Y and is the spatial alignment

- Create an axis/line using TB’s 2 to 1 (+X)

  - This is the planar alignment

- TB 2 is a single point that sets zero in X, Y, Z.

BFW Tooling Ball CSY (BFWTB$CSY)

- Create a plane using Tooling Balls 5, 6, & 7

  - This sets Zero in Y and is the spatial alignment

- Create an axis/line using TB’s 6 to 5 (+X)

  - This is the planar alignment

- TB 6 is a single point that sets zero in X, Y, Z.

| View from Z+ or Downstream | View from X+ Side |
| --- | --- |

Wire CSY (WIRE$CSY)

- Back Plane (-Z) Side sets Zero in Z and is the Spatial Alignment

- Looking from +Z Side, Bottom of Card Sets Zero is the Planar Alignment

- The Intersection of the Left Side of the Card and the Bottom, sets Zero in X & Y

Wire offset from Edges per T. Nakashima’s Optical Measurement

| Wire | Wire Location | Wire Diameter |
| --- | --- | --- |
| Vertical Wire (X Location from Card) | 39.6200 | 0.0383 |
| Horizontal Wire (Y Location from Card) | 37.0187 | 0.0383 |

Support Stand CSY (BASE$CSY)

- A plane on Torque Fixture sets Zero in Y and is the Spatial Alignment

- The symmetry of the outer surfaces where “X” fine adjustments are sets zero in X.

- A line where the Torque Fixture 2X Dowel pins sets zero in Z and is the Planar Alignment.

Beamline CSY (BEAMLINE$CSY)

- Use the BENCH$CSY setup except use the Torque Fixture plate to set zero in Y and the Spatial Alignment, transform +258.67 in Y to replicate Theoretical Beamline Center

- This coordinate system will tell us where the Wire Crosshairs are to a theoretical Beam Centerline.

- The wires should be offset ~ -0.400 in X, -0.200 in Y.

Statistical Data (mm.)

Out of Vacuum

10 Cycles (10 recordings)

Card Locations  
CSY = BENCH$CSY

| STATS | X-POS CARD | X-INTPT CARD | Y-INTPT CARD |
| --- | --- | --- | --- |
| AVG. | 0.07739 | -39.94117 | -95.25854 |
| Range | 0.02047 | 0.02636 | 0.00557 |
| St. Dev. | 0.00590 | 0.00772 | 0.00157 |

Wire Location

CSY = BEAMLINE$CSY

| Stats | X | Y |
| --- | --- | --- |
| AVG. | -0.34277 | -0.22735 |
| Range | 0.02046 | 0.00121 |
| St. Dev. | 0.00591 | 0.00032 |
| Nom. | -0.40000 | -0.20000 |
| RMS Tol | 0.06500 | 0.03000 |

Transformation WIRE$CSY TO BENCH$CSY

| Stats | X | Y | Z | APA_PXPY | PA_PYPZ | PA_PXPZ |
| --- | --- | --- | --- | --- | --- | --- |
| AVG. | -0.34277 | -58.22735 | -2.88558 | 0.03471 | 89.88705 | 90.11467 |
| Range | 0.02046 | 0.00121 | 0.00266 | 0.00803 | 0.00150 | 0.00192 |
| St. Dev. | 0.00591 | 0.00032 | 0.00077 | 0.00241 | 0.00042 | 0.00054 |
| RMS Tol | 0.06500 | 0.03000 | - | - | - | - |

Transformation WIRE$CSY to BFWTB$CSY

| Stats | X | Y | Z | APA_PXPY | PA_PYPZ | PA_PXPZ |
| --- | --- | --- | --- | --- | --- | --- |
| AVG. | 94.39187 | -160.27029 | 35.83288 | 0.19122 | 89.78421 | 89.99193 |
| Range | 0.00139 | 0.00074 | 0.00097 | 0.00246 | 0.00054 | 0.00110 |
| St. Dev. | 0.00044 | 0.00021 | 0.00034 | 0.00084 | 0.00014 | 0.00032 |
| RMS Tol | 0.06500 | 0.03000 | - | - | - | - |

Under Vacuum

10 Cycles (10 recordings)

TB 5-8 Locations

CSY = BENCH$CSY

| TB 5 | X | Y | Z |
| --- | --- | --- | --- |
| AVG. | 94.61579 | 101.79971 | -38.35171 |
| Range | 0.00244 | 0.00406 | 0.00114 |
| St. Dev. | 0.00080 | 0.00157 | 0.00035 |
| RMS Tol | 0.06500 | 0.03000 | - |

| TB 6 | X | Y | Z |
| --- | --- | --- | --- |
| AVG. | -94.26843 | 102.38886 | -38.42323 |
| Range | 0.00206 | 0.00338 | 0.00104 |
| St. Dev. | 0.00070 | 0.00115 | 0.00029 |
| RMS Tol | 0.06500 | 0.03000 | - |

| TB 7 | X | Y | Z |
| --- | --- | --- | --- |
| AVG. | -51.17160 | 102.41353 | 67.42570 |
| Range | 0.00208 | 0.00252 | 0.00119 |
| St. Dev. | 0.00070 | 0.00081 | 0.00038 |
| RMS Tol | 0.06500 | 0.03000 | - |

| TB 8 | X | Y | Z |
| --- | --- | --- | --- |
| AVG. | 51.38056 | 102.10953 | 67.47263 |
| Range | 0.00217 | 0.00290 | 0.00065 |
| St. Dev. | 0.00073 | 0.00094 | 0.00020 |
| RMS Tol | 0.06500 | 0.03000 | - |

Wire Location

CSY = BEAMLINE$CSY

| Stats | X | Y |
| --- | --- | --- |
| AVG. | -0.37798 | -0.21209 |
| Range | 0.00821 | 0.00156 |
| St. Dev. | 0.00294 | 0.00045 |
| Nom. | -0.40000 | -0.20000 |
| RMS Tol | 0.06500 | 0.03000 |

Transformation WIRE$CSY TO BENCH$CSY

| Stats | X | Y | Z | APA_PXPY | PA_PYPZ | PA_PXPZ |
| --- | --- | --- | --- | --- | --- | --- |
| AVG. | -0.39091 | -58.12110 | -2.31314 | 0.01147 | 89.69780 | 90.01288 |
| Range | 0.00821 | 0.00156 | 0.00179 | 0.00220 | 0.00082 | 0.00027 |
| St. Dev. | 0.00294 | 0.00045 | 0.00068 | 0.00080 | 0.00026 | 0.00007 |
| RMS Tol | 0.06500 | 0.03000 | - | - | - | - |

Locations of ¼-20 Jo-Plugs Cylinders Intersected into Plane projected 12.7mm away from surface replicating a ½” (12.7mm) offset to center of Tooling Ball.

CSY = BEAMLINE$CSY

| INTPT | X | Y | Z |
| --- | --- | --- | --- |
| 1 | -26.95171 | 227.65520 | 68.09015 |
| 2 | 27.05826 | 227.64946 | 68.02622 |
| 3 | 79.67686 | 180.00752 | -68.97018 |
| 4 | -79.96459 | 180.02879 | -68.77944 |

# Text boxes

Y+

+Z

Z+

Y+

+X

+Y

X+

Transformation