LCLS BX3 DIPOLE MAGNET

FIDUCIALIZATION REPORT

| Inspector: | Ed Reese |
| --- | --- |
| Responsible Engineer: | Carl Rago |
| Date: | Wednesday, February 14, 2007 |
| Work Order/Charge No.: | 21698-1 |
| Serial Number | 000240 |
| URL of Fiducial Report: | \\Web002\www-group\met\Quality\FIDUCIAL REPORTS\LCLS BX3 DIPOLE\000240.PDF |

Part Set-up – Coordinate System Set-up

Planar Alignment

- Symmetry Axis between inner side plates of housing

Spatial Alignment

- Symmetry Plane between upper and lower poles

“Z” Zero

- Symmetry Point between faces perpendicular to beam

“X” Zero

- Symmetry Axis between Inner side plate of housing

“Y” Zero

- Symmetry Plane between upper and lower poles

Tooling Ball Measurements/Locations

Top Tooling Ball Sockets

| Tooling Ball | FORM | DIAMETER | X | Y | Z |
| --- | --- | --- | --- | --- | --- |
| TB 1 | 0.00024 | 0.49894 | 4.32125 | 4.64740 | -45.19329 |
| TB 2 | 0.00041 | 0.49850 | -4.32259 | 4.64943 | -45.14145 |
| TB 3 | 0.00022 | 0.49787 | 4.31512 | 4.64952 | 45.12009 |
| TB 4 | 0.00025 | 0.49711 | -4.30754 | 4.64828 | 45.10835 |

Side Optical/Magnetic Precision Sphere Holder

| Tooling Ball | FORM | DIAMETER | X | Y | Z |
| --- | --- | --- | --- | --- | --- |
| A | 0.00006 | 1.50006 | 6.62948 | 2.13387 | -31.38417 |
| B | 0.00004 | 1.50006 | 6.62707 | 2.13769 | 31.32683 |
| C | 0.00007 | 1.50007 | 6.61706 | -2.09089 | -31.40045 |
| D | 0.00004 | 1.50008 | 6.61951 | -2.07168 | 31.31292 |

Additional Requested Measurements

| LOCATION | POLE DISTANCE |
| --- | --- |
| -X/-Z | 0.90673 |
| +X/-Z | 0.90626 |
| -X/+Z | 0.90635 |
| +X/+Z | 0.90610 |