LCLS Dipole Magnet

FIDUCIALIZATION REPORT

| Inspector: | Ed Reese |
| --- | --- |
| Responsible Engineer: | Carl Rago |
| Date: | Monday, February 12, 2007 |
| Work Order/Charge No.: | 21697-1 |
| Serial Number | 000354 |
| URL of Fiducial Report: | \\Web002\www-group\met\Quality\FIDUCIAL REPORTS\LCLS DIPOLE\000354.PDF |

Part Set-up – Coordinate System Set-up

Planar Alignment

- Symmetry Axis between inner side plates of housing

Spatial Alignment

- Symmetry Plane between upper and lower poles

“Z” Zero

- Symmetry Point between faces perpendicular to beam

“X” Zero

- Symmetry Axis between Inner side plate of housing

“Y” Zero

- Symmetry Plane between upper and lower poles

-Z End View

Tooling Ball Measurements/Locations

| Tooling Ball | FORM | DIAMETER | X | Y | Z |
| --- | --- | --- | --- | --- | --- |
| TB 1 | 0.00002 | 0.50000 | 3.87609 | 5.16472 | -17.03492 |
| TB 2 | 0.00046 | 0.49832 | -3.85739 | 5.16609 | -16.97847 |
| TB 3 | 0.00013 | 0.49758 | 3.92304 | 5.16869 | 16.97503 |
| TB 4 | 0.00011 | 0.49921 | -3.86078 | 5.16535 | 16.99323 |

Additional Requested Measurements

| LOCATION | POLE DISTANCE |
| --- | --- |
| -X/-Z | 1.37563 |
| +X/-Z | 1.37646 |
| -X/+Z | 1.37886 |
| +X/+Z | 1.38027 |

# Text boxes

TB1

TB2

TB3

+X

+Y

TB4