LCLS Injector Dogleg Dipole Magnet

FIDUCIALIZATION REPORT

| Inspector: | Keith Caban |
| --- | --- |
| Responsible Engineer: | Roger Carr |
| Date: | Monday, May 15, 2006 |
| Work Order/Charge No.: | 92-4217-3 |
| Serial Number | 000437 |

Part Set-up – Coordinate System Set-up

Planar Alignment

- Mid-Plane of the magnet

Spatial Alignment

- Bottom pole plane

  - -Y plane

“Z” Zero

- Mid-Plane of the magnet

“X” & “Y” Zero

- X zero is created by the symmetry point of the inner stainless steel where coils lie.

    - Both ends of both sides.

      - Create lines which is y axis.

        - This Creates X- zero

- Symmetry of pole planes

  - This creates Y- zero

| ISO | BACK VIEW |
| --- | --- |

Tooling Ball Measurements/Locations

Top of magnet; view from “+Y”

| Tooling Ball | FORM | DIAMETER | X | Y | Z |
| --- | --- | --- | --- | --- | --- |
| TB 1 | 0.00094 | 0.49676 | 7.31527 | 7.77509 | 2.90755 |
| TB 2**** | 0.00177 | 0.48952 | 7.30949 | 7.77900 | -2.89354 |
| TB 3 | 0.00073 | 0.49641 | -7.32221 | 7.78232 | -2.89451 |
| TB 4^ | 0.00125 | 0.49310 | -7.32441 | 7.78207 | 2.91233 |

**** - denotes bad tooling ball socket holder, part was out of round and undersized had to use a smaller shaft 0.2475 because standard did not fit hole.

**** Y value changed 7/28/06 reamed out hole to size.

^ - denotes that tooling ball socket holder was oversized more than normal see form and diameter size.

Additional Requested Measurements

Parallelism of pole planes

// = 0.00088

# Text boxes

TB 1

TB 2

TB 3

TB 4

Z+

+Z

+X

A

B

C

D

View From +Z

+Y

+Z

+Y

+X

+X

+Z