LCLS ‘Q150kG’ LTU QUADRUPOLE MAGNET

FIDUCIALIZATION REPORT

| Inspector: | Keith Caban |
| --- | --- |
| Responsible Engineer: | Carl Rago |
| Date: | Wednesday, December 05, 2007 |
| Work Order/Charge No.: | 9242609 |
| Serial Number: | SLAC – 002013 / SN 014 |
| URL of Fiducial Report: | \\Web002\www-group\met\Quality\FIDUCIAL REPORTS\LCLS LTU Q150kG QUADS\002013.pdf |

Part Set-up – Coordinate System Set-up

Spatial Alignment

- Geometric axis of the poles of the magnet.

Planar Alignment

- Plane on top of magnet where tooling ball sockets are welded to.

“Z” Zero

- Mid-plane of the magnet (middle of upstream and downstream ends).

“X” & “Y” Zero

- Geometric axis of the poles of the magnet.

| Front View | Side View |
| --- | --- |

Tooling Ball Measurements/Locations

Tooling Ball Adapter Cylinder Projected 1” Offset to the Tooling Ball Adapter Plane

| Tooling Ball | Cyl. Dia. | X | Y | Z |
| --- | --- | --- | --- | --- |
| Proj. TB 1 | 0.25005 | 2.66651 | 10.21003 | 5.60837 |
| Proj. TB 2 | 0.24992 | 2.66788 | 10.20399 | -5.60948 |
| Proj. TB 3 | 0.25009 | -2.64442 | 10.20704 | -5.61091 |
| Proj. TB 4 | 0.24999 | -2.64199 | 10.21256 | 5.60608 |

Pole Data

*Data looking from Downstream End

| Side | Pole Dist A-C | Pole Dist B-D | Gap 1 | Gap 2 | Gap 3 | Gap 4 |
| --- | --- | --- | --- | --- | --- | --- |
| Downstream End (+Z) | 1.26063 | 1.26113 | 0.43383 | 0.43667 | 0.43336 | 0.43231 |
| Upstream End (-Z) | 1.25891 | 1.25888 | 0.43302 | 0.43345 | 0.43074 | 0.43088 |

# Text boxes

+Y

+Z

+Y

+X

+Z

+X