LCLS ‘0.91Q17.72’ LTU Quadrupole Magnet

FIDUCIALIZATION REPORT

| Inspector: | Keith Caban |
| --- | --- |
| Responsible Engineer: | Carl Rago |
| Date: | Monday, March 26, 2007 |
| Work Order/Charge No.: | 21699-1 |
| Serial Number: | 000311 |
| URL of Fiducial Report: | \\Web002\www-group\met\Quality\FIDUCIAL REPORTS\LCLS LTU QUAD\000311.pdf |

Part Set-up – Coordinate System Set-up

Spatial Alignment

- Geometric axis of the poles of the magnet.

Planar Alignment

- Clocking plane where TB A,C, D, & E reside.

“Z” Zero

- Mid-plane of the magnet (pole planes on each end).

“X” & “Y” Zero

- Geometric axis of the poles of the magnet.

| Front View | Side View |
| --- | --- |

Tooling Ball Measurements/Locations

Top of magnet; view from “+Y”

| Tooling Ball | FORM | DIAMETER | X | Y | Z |
| --- | --- | --- | --- | --- | --- |
| TB A | 0.00008 | 0.49945 | 5.02994 | 3.70497 | -8.07217 |
| TB B | 0.00035 | 0.49766 | -1.33017 | 7.32687 | -7.97999 |
| TB C | 0.00014 | 0.49953 | -5.02533 | 3.69999 | -8.04243 |
| TB D | 0.00004 | 0.49943 | -5.26053 | 3.70315 | 6.27197 |
| TB E | 0.00014 | 0.49935 | -1.37415 | 7.32568 | 7.95803 |
| TB F | 0.00037 | 0.49833 | 5.25181 | 3.70714 | 6.30125 |

.

Pole Distances

| Pole | -Z side | +Z side |  |
| --- | --- | --- | --- |
| A-C | 0.90596 | 0.90630 | 0.00034 |
| B-D | 0.90586 | 0.90701 | 0.00115 |

# Text boxes

+Y

+Z

+Y

+X

+X

+Z

D

C

B

A