LCLS ‘0.91Q17.72’ LTU Quadrupole Magnet

FIDUCIALIZATION REPORT

| Inspector: | Keith Caban |
| --- | --- |
| Responsible Engineer: | Carl Rago |
| Date: | Wednesday, February 28, 2007 |
| Work Order/Charge No.: | 21699-1 |
| Serial Number: | 000314 |
| URL of Fiducial Report: | \\Web002\www-group\met\Quality\FIDUCIAL REPORTS\LCLS LTU QUAD\000314.pdf |

Part Set-up – Coordinate System Set-up

Spatial Alignment

- Geometric axis of the poles of the magnet.

Planar Alignment

- Clocking plane where TB A,C, D, & E reside.

“Z” Zero

- Mid-plane of the magnet (pole planes on each end).

“X” & “Y” Zero

- Geometric axis of the poles of the magnet.

| Front View | Side View |
| --- | --- |

Tooling Ball Measurements/Locations

Top of magnet; view from “+Y”

| Tooling Ball | FORM | DIAMETER | X | Y | Z |
| --- | --- | --- | --- | --- | --- |
| TB A | 0.00019 | 0.49742 | 5.05497 | 3.71420 | -8.09705 |
| TB B | 0.00035 | 0.49833 | -1.45509 | 7.31688 | -8.05082 |
| TB C | 0.00009 | 0.49973 | -5.04517 | 3.69186 | -8.05783 |
| TB D | 0.00052 | 0.49808 | -5.30611 | 3.69181 | 6.27965 |
| TB E | 0.00003 | 0.49996 | -1.43912 | 7.31738 | 8.06021 |
| TB F | 0.00004 | 0.49837 | 5.30018 | 3.71439 | 6.30034 |

.

Pole Distances

| Pole | -Z side | +Z side |  |
| --- | --- | --- | --- |
| A-C | 0.90711 | 0.90685 | 0.00026 |
| B-D | 0.90693 | 0.90614 | 0.00079 |

# Text boxes

+Y

+Z

+Y

+X

+X

+Z

D

C

B

A