LCLS Laser Heater Chicane Dipole Magnet

FIDUCIALIZATION REPORT

| Inspector: | Keith Caban |
| --- | --- |
| Responsible Engineer: | C. Rago |
| Date: | Monday, December 17, 2007 |
| Work Order/Charge No.: | 96-8337-5 |
| Serial Number: | LCLS - 002008 MAG SN – 42611-1 |
| URL of Fiducial Report: | \\Web002\www-group\met\Quality\FIDUCIAL REPORTS\LCLS Laser Heater Dipole Magnets\002008.pdf |

Part Set-up – Coordinate System Set-up

** Poles have been modified

Planar Alignment

- Mid-Plane of the magnet

Spatial Alignment

  - Plane along the bottom 2 planes.

“Z” Zero

- Mid-Plane of the magnet. (see dashed planar alignment above)

“X” & “Y” Zero

- X zero is created by the symmetry point of the inner stainless steel where coils lie.

    - Both ends of both sides.

      - Create lines which is y axis.

        - This Creates X- zero

- Symmetry of pole planes

  - This creates Y- zero

Tooling Ball Measurements/Locations

Projected cylinder into 1” offset plane

| Tooling Ball | FORM | DIAMETER | X | Y | Z | b |
| --- | --- | --- | --- | --- | --- | --- |
| TB 1 | 0.00173 | 0.25066 | 5.20207 | 6.50395 | 1.06838 | 0.00295 |
| TB 2 | 0.00230 | 0.25052 | 5.20016 | 6.50654 | -1.05206 | 0.00156 |
| TB 3 | 0.00100 | 0.24998 | -5.19613 | 6.50961 | -1.05620 | 0.00089 |
| TB 4 | 0.00082 | 0.25004 | -5.19268 | 6.50711 | 1.07358 | 0.00181 |

Pole Gaps

| GAP | GAP DIST | STEP | STEP DIST | LG GAP | LG DIST |
| --- | --- | --- | --- | --- | --- |
| 14P | 1.18091 | P1 | 0.00943 | P | 1.19751 |
| 23P | 1.18131 | P2 | 0.00703 | N | 1.19782 |
| 14N | 1.18090 | N1 | 0.00942 |  |  |
| 23P | 1.18107 | N2 | 0.00736 |  |  |

# Text boxes

+Z

+X

+Y

+Y

+Z

+X