LCLS Laser Heater Chicane Dipole Magnet

FIDUCIALIZATION REPORT

| Inspector: | Keith Caban |
| --- | --- |
| Responsible Engineer: | C. Rago |
| Date: | Tuesday, December 18, 2007 |
| Work Order/Charge No.: | 96-8337-5 |
| Serial Number: | LCLS - 002009 MAG SN – 42611-2 |
| URL of Fiducial Report: | \\Web002\www-group\met\Quality\FIDUCIAL REPORTS\LCLS Laser Heater Dipole Magnets\002009.pdf |

Part Set-up – Coordinate System Set-up

Planar Alignment

- Mid-Plane of the magnet

Spatial Alignment

  - Plane along the bottom 2 planes.

“Z” Zero

- Mid-Plane of the magnet. (see dashed planar alignment above)

“X” & “Y” Zero

- X zero is created by the symmetry point of the inner stainless steel where coils lie.

    - Both ends of both sides, lower pole only.

      - Create lines which is y axis.

        - This Creates X- zero

- Symmetry of pole planes

  - This creates Y- zero

Tooling Ball Measurements/Locations

Projected cylinder into 1” offset plane

| Tooling Ball | FORM | DIAMETER | X | Y | Z | b |
| --- | --- | --- | --- | --- | --- | --- |
| TB 1 | 0.00120 | 0.25109 | 5.19374 | 6.51232 | 1.07206 | 0.00166 |
| TB 2 | 0.00088 | 0.25067 | 5.19674 | 6.51306 | -1.05544 | 0.00093 |
| TB 3 | 0.00114 | 0.25077 | -5.19546 | 6.51388 | -1.06213 | 0.00059 |
| TB 4 | 0.00244 | 0.25193 | -5.19697 | 6.51299 | 1.06379 | 0.00067 |

Pole Gaps

| GAP | GAP DIST | STEP | STEP DIST | LG GAP | LG DIST |
| --- | --- | --- | --- | --- | --- |
| 14P | 1.18040 | P1 | 0.00734 | P | 1.19594 |
| 23P | 1.18091 | P2 | 0.00795 | N | 1.19579 |
| 14N | 1.18034 | N1 | 0.00706 |  |  |
| 23P | 1.18095 | N2 | 0.00808 |  |  |

# Text boxes

+Z

+X

+Y

+Y

+Z

+X