LCLS Laser Heater Chicane Dipole Magnet

FIDUCIALIZATION REPORT

| Inspector: | Keith Caban |
| --- | --- |
| Responsible Engineer: | C. Rago |
| Date: | Tuesday, December 18, 2007 |
| Work Order/Charge No.: | 96-8337-5 |
| Serial Number: | LCLS - 002010 MAG SN – 42611-4 |
| URL of Fiducial Report: | \\Web002\www-group\met\Quality\FIDUCIAL REPORTS\LCLS Laser Heater Dipole Magnets\002010.pdf |

Part Set-up – Coordinate System Set-up

Planar Alignment

- Mid-Plane of the magnet

Spatial Alignment

  - Plane along the bottom 2 planes.

“Z” Zero

- Mid-Plane of the magnet. (see dashed planar alignment above)

“X” & “Y” Zero

- X zero is created by the symmetry point of the inner stainless steel where coils lie.

    - Both ends of both sides, lower pole only.

      - Create lines which is y axis.

        - This Creates X- zero

- Symmetry of pole planes

  - This creates Y- zero

Tooling Ball Measurements/Locations

Projected cylinder into 1” offset plane

| Tooling Ball | FORM | DIAMETER | X | Y | Z | b |
| --- | --- | --- | --- | --- | --- | --- |
| TB 1 | 0.00087 | 0.25022 | 5.19466 | 6.50845 | 1.06585 | 0.00108 |
| TB 2 | 0.00093 | 0.25076 | 5.19361 | 6.50942 | -1.05899 | 0.00148 |
| TB 3 | 0.00082 | 0.25058 | -5.19853 | 6.50877 | -1.06150 | 0.00081 |
| TB 4 | 0.00208 | 0.25190 | -5.20092 | 6.50778 | 1.06280 | 0.00228 |

Pole Gaps

| GAP | GAP DIST | STEP | STEP DIST | LG GAP | LG DIST |
| --- | --- | --- | --- | --- | --- |
| 14P | 1.18120 | P1 | 0.00939 | P | 1.19762 |
| 23P | 1.18138 | P2 | 0.00699 | N | 1.19762 |
| 14N | 1.18112 | N1 | 0.00918 |  |  |
| 23P | 1.18126 | N2 | 0.00720 |  |  |

# Text boxes

+Z

+X

+Y

+Y

+Z

+X