LCLS ‘QE’ QUADRUPOLE MAGNET

FIDUCIALIZATION REPORT

| Inspector: | Keith Caban |
| --- | --- |
| Responsible Engineer: | Carl Rago |
| Date: | Friday, October 26, 2007 |
| Work Order/Charge No.: | 23188-49 |
| Serial Number: | 002798 |
| URL of Fiducial Report: | \\Web002\www-group\met\Quality\FIDUCIAL REPORTS\LCLS QE QUAD'S\002798.pdf |

Part Set-up – Coordinate System Set-up

Spatial Alignment

- Geometric axis of the poles of the magnet.

Planar Alignment

- A 4 point plane on the top of where 2 arrows indicate in the picture below.

“Z” Zero

- Mid-Plane of the magnet

“X” & “Y” Zero

- Geometric Axis of the Poles of the magnet.

| Side View | Front View |
| --- | --- |

Tooling Ball Measurements/Locations

Top of magnet; view from “+Y”

| Tooling Ball | FORM | DIAMETER | X | Y | Z |
| --- | --- | --- | --- | --- | --- |
| TB 1 | 0.00002 | 0.50004 | 8.24443 | 1.21314 | 1.75578 |
| TB 2 | 0.00010 | 0.49915 | 8.25298 | 1.19965 | -1.22083 |
| TB 3 | 0.00009 | 0.49922 | 1.16572 | 8.27989 | 1.70096 |
| TB 4 | 0.00017 | 0.49907 | 1.19201 | 8.26105 | -1.27926 |
| TB 5 | 0.00012 | 0.49935 | -8.26094 | 1.20963 | 1.70687 |
| TB 6 | 0.00003 | 0.49948 | -8.25074 | 1.20051 | -1.26231 |

Pole Distances

| Pole | +Z side | -Z side |  |
| --- | --- | --- | --- |
| A-C | 1.08574 | 1.08668 | 0.00094 |
| B-D | 1.08502 | 1.08729 | 0.00227 |

*** Points collected approx. 0.125 from each end

*** Picture to the left is looking down from +Z

``

# Text boxes

TB 6

TB 5

TB 2

TB 1

+Y

+Z

+Y

+X

+X

+Z

D

C

B

A