LCLS Gun Spectrometer Quad Dipole Corrector Magnet

FIDUCIALIZATION REPORT

| Inspector: | Keith Caban |
| --- | --- |
| Responsible Engineer: | Roger Carr |
| Date: | Monday, May 15, 2006 |
| Work Order/Charge No.: | 20966-1 Task – 30 |
| Serial Number | 1 |

Part Set-up – Coordinate System Set-up

Planar Alignment

- Mid-Plane of the magnet

Spatial Alignment

- A line on the top part of the magnet

  - +X is from TB 1 & 2 side to TB 3 & 4 Side

“Z” Zero

- Mid-Plane of the magnet

“X” & “Y” Zero

- On both ends

  - Tangent point of each radii (4 on each end, 8 total).

    - Create a line between diagonal tangent points creates 2 lines.

      - Intersect the lines.

        - Creates a point on each end.

- Create a line of these 2 end points

  - This is the “X” & “Y” Zero, and Beamline or” Z” Axis.

| ISO | BACK VIEW |
| --- | --- |

Tooling Ball Measurements/Locations

Top of magnet; view from “+Y”

| Tooling Ball | FORM | DIAMETER | X | Y | Z |
| --- | --- | --- | --- | --- | --- |
| TB 1 | 0.00065 | 0.49657 | -3.51001 | 4.82224 | -0.68423 |
| TB 2 | 0.00090 | 0.49673 | -3.51151 | 4.82349 | 0.66619 |
| TB 3 | 0.00073 | 0.49698 | 3.54427 | 4.82099 | -0.68228 |
| TB 4 | 0.00097 | 0.49708 | 3.54301 | 4.82216 | 0.66856 |

Additional Requested Measurements

Distance from Tangent point of 4X poletips

A-C = 1.66521

B-D = 1.66543

Parallelism of 4X pole tips to the beamline

- 0.00037

- 0.00033

- 0.00044

- 0.00052

# Text boxes

TB 1

TB 2

TB 3

TB 4

Z+

+Z

+X

A

B

C

D

View From +Z

+Y

-Z

+Y

+X

+X

+Z