LCLS Injector Quadrupole Magnet

FIDUCIALIZATION REPORT

| Inspector: | Keith Caban |
| --- | --- |
| Responsible Engineer: | Roger Carr |
| Date: | Friday, April 28, 2006 |
| Work Order/Charge No.: | N/A |
| Serial Number | 000381 |

Part Set-up – Coordinate System Set-up

Planar Alignment

- Mid-Plane of the magnet

Spatial Alignment

- A line on the top part of the magnet

  - +X goes towards (Magnet Info Label)

“Z” Zero

- Mid-Plane of the magnet

“X” & “Y” Zero

- On both ends

  - Tangent point of each radii (4 on each end, 8 total).

    - Create a line between diagonal tangent points creates 2 lines.

      - Intersect the lines.

        - Creates a point on each end.

- Create a line of these 2 end points

  - This is the “X” & “Y” Zero, and Beamline or” Z” Axis.

Tooling Ball Measurements/Locations

Top of magnet; view from “+Y”

| Tooling Ball | FORM | DIAMETER | X | Y | Z |
| --- | --- | --- | --- | --- | --- |
| TB 1 | 0.00087 | 0.49507 | 6.49838 | 9.37867 | -1.25062 |
| TB 2 | 0.00052 | 0.49718 | -6.49606 | 9.37673 | -1.25002 |
| TB 3 | 0.00063 | 0.49606 | -6.49845 | 9.37954 | 1.24732 |
| TB 4 | 0.00065 | 0.49661 | 6.5000 | 9.37660 | 1.24801 |

Additional Requested Measurements

Tangent Point Straightness

- 0.00021

- 0.00026

- 0.00020

- 0.00024

Parallelism to Beamline

- 0.00030

- 0.00017

- 0.00022

- 0.00046

Distance of opposite Tangent

Axis

A-C = 1.25874 (0.62936, 0.62995)

B-D = 1.25986 (0.62991, 0.62938)

# Text boxes

TB 1

TB 2

TB 3

TB 4

Z+

+Z

+X

A

B

C

D

View From +Z