LCLS Injector Quadrupole Magnet

FIDUCIALIZATION REPORT

| Inspector: | Keith Caban |
| --- | --- |
| Responsible Engineer: | Roger Carr |
| Date: | Thursday, August 17, 2006 |
| Work Order/Charge No.: | 92-4215-8 |
| Serial Number | 000383 |

Part Set-up – Coordinate System Set-up

Planar Alignment

- Mid-Plane of the magnet

Spatial Alignment

- A line on the top part of the magnet

  - +X goes towards (Magnet Info Label)

“Z” Zero

- Mid-Plane of the magnet

“X” & “Y” Zero

- On both ends

  - Tangent point of each radii (4 on each end, 8 total).

    - Create a line between diagonal tangent points creates 2 lines.

      - Intersect the lines.

        - Creates a point on each end.

- Create a line of these 2 end points

  - This is the “X” & “Y” Zero, and Beamline or” Z” Axis.

Tooling Ball Measurements/Locations

Top of magnet; view from “+Y”

| Tooling Ball | FORM | DIAMETER | X | Y | Z |
| --- | --- | --- | --- | --- | --- |
| TB 1 | 0.00096 | 0.49535 | 6.50065 | 8.87767 | -1.25051 |
| TB 2 | 0.00049 | 0.49793 | -6.49896 | 8.87678 | -1.24943 |
| TB 3 | 0.00032 | 0.49870 | -6.49858 | 8.88167 | 1.25141 |
| TB 4 | 0.00062 | 0.49804 | 6.50165 | 8.87745 | 1.24871 |

Additional Requested Measurements

Tangent Point Straightness

- 0.00038

- 0.00027

- 0.00053

- 0.00037

Parallelism to Beamline

- 0.00045

- 0.00066

- 0.00057

- 0.00046

Distance of opposite Tangent

Axis

A-C = 1.26003 (0.62994, 0.63009)

B-D = 1.25919 (0.62954, 0.62964)

# Text boxes

TB 1

TB 2

TB 3

TB 4

Z+

+Z

+X

A

B

C

D

View From +Z